Signal Generator

SG62

Instruction Manual

ADVANCE COMPONENTS LIMITED

INSTRUMENT DIVISION

ROEBUCK ROAD, HAINULT, ILFORD, ESSEX, ENGLAND

TELEPHONE: HAINULT 4444 TELEGRAMS: ATTENUATE ILFORD

Publication 135

Contents

Section 1	Introduction							4

Section 2	Specification							5

Section 3	Operation							7

3.1	Supply Voltage						7

3.2	Frequency Selection						7

3.3	R.F. Output							7

3.4	Output Termination						7

3.5	Modulation and CW						7

3.6	Audio Frequency Output					7

3.7	Connections to Points of High Potential			8

Section 4	Maintenance							9

4.1	Removal of Instrument Case					9

4.2	Replacement of Instrument Case				9

4.3	Valve Replacement						9

Section 5	Components List						10

Section 6	Factory Service						13

Illustration

Fig. 1	Circuit Diagram						12

3

1

1. Introduction

The Wide Band Signal Generator SG62 provides r.f. signals over the frequency range 150 kc/s to 220 Mc/s, thus covering most radio and television bands.  The instrument employs a Colpitts oscillator with six ranges that are displayed on a calibrated dial having a total length of 50 inches.  A band spreading device maintains an open scale at the high frequency end of each range.

An overall frequency accuracy of  ± 1% is achieved and an adjustable cursor provides means of obtaining greater accuracy when checked against standard frequencies, such as a broadcast carrier signal or crystal calibrator.

The output signal is available as a continuous sine wave or may be amplitude modulated 30% at 400 c/s, when an audio frequency signal is also available at separate terminals.  The output level is controlled by an accurate attenuator system  which permits repeatable measurements with signal levels as low as 1µV.

4

2

2   Specification

Frequency Ranges	150 kc/s – 500 kc/s

500 kc/s – 1.m Mc/s

1.8 Mc/s – 6 Mc/s

6 Mc/s – 20 Mc/s

20 Mc/s – 70 Mc/s

70 Mc/s – 220 Mc/s

Accuracy ± 1% with cursor in central position.

R/F. Output		The output is fed through a 75Ω cable, which should be                        			terminated either by termination unit, type TP2B, or any other                      			load of 75Ω. The voltage into the terminating load is adjustable                                             			from 1µV to 100 mV by means of a five position attenuator and                                    			a continuously variable control.

Accuracy of maximum output ±3db.

Accuracy of step attenuator ± (3db+3 µV).

Output Impedance	The output impedance at the end of the output cable is 75Ω                     			when it is not terminated, except at the x10mV position of the                                 			step attenuator when it varies  from 60 to 85Ω. When                         			terminated the impedance becomes 37.5Ω.

Modulation		The r.f. signal is available as either a continuous                        		     	wave or midulated 30% ± 3db at  400 c/s ±10%.

Audio Frequency	An audio frequency signal of 400 c/s at approximately 10V is                                      			available into a high impedance when the oscillator is                			modulated.

Valves and		1 Valve type 12AT7

Accessories		1 Neon Pilot Lamp, 100-125V.

1 75Ω Termination Pad type TP2B.

1 Connector Cable R.F. , type PL28.

Power Supply		The standard model is operated from power supplies of 105-           			125V, 140-160V, 210-250V, 40-100c/s.  Consumption                 			approximately 20W.

Weight	10 lb (4.5kg).

5

2

Dimensions		12 in. X 9 in. X 6 in.  (30.5 cm, X 22.9 cm, X 15.2 cm).

Finish			Dark blue metal case complete with leather carrying handle..                      			Light grey front panel with medium grey surround.  Knobs            			medium grey.  Colours to B.S. 2660.  Case tint No. 7-086, front             			panel tint No. 9-093, front panel surround tint No. 9-095.

6

3

# 3   Operation

3.1	Supply Voltage

When the instrument is first received it should be checked to en-sure that it is set to the correct supply voltage. First remove the instrument case, para. 4.1. Adjust the supply transformer voltage tap to the nominal voltage of the local supply and replace the case, para. 4.2.

3.2	Frequency Selection

A signal frequency in the range from 150 kc/s to 220 Mc/s may be selected using the range switch and the calibrated scale. The slow motion drive facilitates fine tuning. Frequency accuracy is ± 1% and for increased accuracy a cursor is provided. This is adjustable over a small angle by means of a screw head on the axis of rotation of the pointer. The scale reading may be adjusted to the exact value by checking the output from the SG62 against the sound carrier of a television signal, or the unmodulated carrier of an F.M. transmission, for example, using a receiver tuned to the correct frequency.

3.3	R.F. Output

The output voltage is obtained at the end of a terminated 75Ω cable. The voltage is controlled by two attenuators marked OUTPUT VOLTAGE, a fine control. calibrated 0-10 and a step attenuator calibrated: Xl, Xl0, Xl00 µV and Xl, X10 mV. By multiplying the indication of the controls the output may be set from 1µV to l00mV.

3.4	Output Termination

The termination pad type TP2B is provided to terminate the r.f. connector cable correctly and to prevent standing waves. This pad incorporates a 75Ω resistor and the output signal may be taken from it to the instrument under test, keeping the leads as short as possible. Alternatively, the output may be fed directly to a load of 75Ω.

3.5	Modulation and CW

When the MOD/CW switch is in the MOD position the r.f. signal is modulated at 400 C/S (±10%), to a depth of 30% (± 3 dB).

When the switch is in the CW position the r.f signal is unmodulated.

3.6	Audio Frequency Output

An output of 400 c/s at approximately l0V into a high impedance is obtainable from the A.F. output socket, when the MOD/CW switch is in the MOD position.

7

3

3.7	Connections to Points of High Potential

As the generator has an output with a low resistance path to earth, it is necessary to use a blocking capacitor (e.g. 0.05µF at 150 kc/s) if a signal is to be injected at a point of high potential.

If the chassis of the instrument under test is 'live', as found on an a.c./d.c. receiver, capacitors should be fitted in both the output and earth connecting leads.

8

4

# 4 Maintenance

4.1	Removal of Instrument Case

(1)	Remove the four screws at the rear.

(2)  Lay the instrument on its back and carefully lift the chassis and front panel clear of the surrounding case.

4.2	Replacement of Instrument Case

When reassembling, ensure that the case is correctly located in the groove at the rear of the front panel surround before replacing the four retaining screws.

4.3	Valve Replacement

The oscillator valve 12AT7 is contained in the internal screening box which may be seen when the instrument case has been removed (Para 4.1). The cover of the screening box may be screwed off, exposing the valve and oscillator assembly.

Changing the oscillator valve should not appreciably affect the frequency calibration, providing that the trimmer and coil assembly are not disturbed.

When replacing the oscillator cover, ensure that it is screwed on firmly without excessive tightening which wilt cause difficulty if the cover is subsequently removed.

9

5

## 5 Components List and Circuit Diagram

### RESISTORS

Ref.	Description			Tol.		Rating		Part No.

R 1	560Ω   Carbon			10%		¼ W		9236

R 2	100KΩ Carbon			10%		¼ W		1270

R 3	27KΩ  Carbon			10%		¼ W		868

R 4	33KΩ	Carbon			10%		¼ W		10476

R 5	22kΩ	Carbon			10%		¼ W		1271

R 6	22kΩ	Carbon			10%	       Morgan “T”	10375

R 7	39Ω	Carbon			10%		½ W		12469

R 8	750Ω	Carbon			5%		¼ W		3386

R 9	91Ω	Carbon			5%		¼ W		3388

R10	750Ω	Carbon			5%		¼ W		3386

R11	91Ω	Carbon			5%		¼ W		3388

R12	750Ω	Carbon			5%		¼ W		3386

R13	91Ω	Carbon			5%		¼ W		3388

R14	750	Carbon			5%		¼ W		3386

R15	82Ω	Carbon			5%		¼ W		3387

R16	Selected on Test

R17	470Ω	Carbon			5%		¼ W		1671

R18	1KΩ	Carbon			5%		¼ W		1175

R19	56KΩ	Carbon			5%		¼ W		4406

R20	1KΩ	Carbon			5%		¼ W		1175

R21	Pot. 100Ω							10374

R22	120KΩ	Carbon			5%		¼ W		4407

R23	47Ω	Carbon			5%		½ W		10917

R24	22Ω	Carbon			5%		½ W		11755

CAPACITORS

C 1	Wire Trimmer							10681

C 2	Wire Trimmer							10681

C 3	Wire Trimmer							10681

C 4	Wire Trimmer							10681

C 4	Wire Trimmer							10681

C 5	Wire Trimmer							10681

C 6	Wire Trimmer							10681

C 7	Var. 532 + 532pf						10555

C 8	50pf Ceramic							3699

C 9	.04µF Hunts W99		20%		150V D.C.	7485

C 10 	300pf Lead Through						7099

C 11	.04µF Hunts W99		20%		150V D.C.	7485

C12	300pf Lead Through						7099

C13	.04µF Hunts W99		20%		150V D.C.	7485

C14	.04µF Hunts W99		20%		150V D.C.	7485

C15	300pf Lead Through						7099

C16	.04µF Hunts W99		20%		150V D.C.	7485

C17	.04µF Hunts W99		20%		150V D.C.	7485

C18	.04µF Hunts W99		20%		150V D.C.	7485

10

#### 5

#### CAPACITORS

Ref.		Description			Tol.		Rating		Part No.

C19		.04µF Hunts W99		20%		150V D.C.	7485

C20+C21  	16 µf +16 µF					350V D.C.	7014

C22		.005µF W99			20%		400V D.C.	8780

C23		.005µF W99			20%		400V D.C.	8780

C24		.04µF Hunts W99		20%		150V D.C.	7485

C25		Wire Trimmer							10681

### INDUCTORS

Ref.	Description					Part No.

L 1	Band “F”   Tuning				RF.594

L1A	Band “F”  Coupling				RF594

L 2	Band “E”  Tuning				RF595

L2A	Band “E”  Coupling				RF595

L 3	Band “D”  Tuning				RF596

L3A	Band “D”  Coupling				RF596

L 4	Band “C”  Tuning				RF597

L4A	Band “C”  Coupling				RF597

L 5	Band “B”  Tuning				RF598

L5A	Band “B”  Coupling				RF598

L 6	Band “A”  Tuning				S/Plate

L 7	Band “A”  Tuning				Strip

L 8	Band “A”  Coupling				Part of Wiring

L 9	Heater Choke					C.95

L10	R.F. Choke					C.16

L11	R.F. Choke					C.16

L12	Heater Choke					C.95

L13	A.C. Supply Filter Choke			C.96

L14	A.C. Supply Filter Choke			C.96

L15	Choke, Dubilier Type 666			11212

### MISCELLANEOUS

T1	Transformer, Modulation			MT.348

T2	Transformer, Power				MT.345/B

MR1	Rectifier					8388

LP1	Pilot Lamp					879

V1	Valve 12AT7					7106

SW1, SW1A, SW1B	Range Switch			10364

SW2,SW2A	Supply Switch-CW/MOD Switch	10376

75Ω Termination Pad					TP2B

R.F. Connector Cable					PL28

Step Attenuator complete assembly (less Knob)	10632

11

5

12

6

##### 6 Factory Service

Our factory Service Department is at your disposal should you wish to obtain further repair information by telephone or letter. The Type and Serial Number of the instrument should always be quoted. We maintain an efficient Service facility and the instrument can, if desired, be returned to our factory for repair.

The instrument is guaranteed for a period of one year from its delivery to the purchaser for the replacement of defective parts, other than valves, semi-conductors and fuses.

Valves and semi-conductors are subject to the manufacturer's guarantee.

Equipment returned to us for servicing must be adequately packed, preferably in the special box supplied, and shipped with transportation charges prepaid. We can accept no responsibility for instruments arriving damaged. Should the cause of failure during the guarantee period be due to misuse or abuse of the instrument, or if the guarantee has expired, the repair will be charged and put in hand without delay unless other instructions are received.

OUR SALES, SERVICE AND ENGINEERING DEPARTMENTS

ARE AT YOUR SERVICE AT ALL TIMES

###### ADVANCE COMPONENTS

ROEBUCK ROAD, HAINULT, ILFORD, ESSEX, ENGLAND

Telephone: HAINULT 4444

Telegrams: ATTENUATE, ILFORD

13