Purpose:  The focus pot used in the video deflection board is under-rated and sometimes burns up.  It must be replaced with a new 1000V pot.  The new pot has a ceramic substrate and a black shaft, the original has a blue adjusting dial.

Parts Required:

182-000-205		qty 1	1000V 2M pot

594-120-001		qty 1	Tie-wrap

Description:

- Remove cover from video deflection board.

- Remove 4 cables from video deflection board.

- Cut tie-wrap securing anode lead wire to yoke tightening arm.

- Remove anode lead from CRT (discharge tube to ground using an insulated screwdriver, tube is generally self-discharging but this step should always be taken as a precaution.)

- Remove video deflection board.

- Remove and replace the focus pot (VR302).

- Re-assemble using steps 6-1 above.

- Secure anode lead to yoke tightening arm using a tie-wrap.