Description:	This procedure will guide you through the steps required to remove the LS-NET option and install the LS-GPIB option.

Warning:  Disassembly described in this section is perfromed with power removed from the instrument.  Do not remove or replace any of the circuit board assemblies without first verifying that power has been removed from the instrument.  Anti-static protection measures should be taken prior to the handling of any printed circuit boards.

Procedure:

1.	Disconnect power cable from the instrument.

2.	Place the LS140 face down into the plastic front panel protective cover that was shipped with the unit or on a soft surface to protect the BNC’s and front panel knob.

3.	Remove the four screws attaching the rear feet to the rear bezel.

4.	Remove the six screws attaching the cabinet to the rear bezel.

5.	Hold the cabinet by the two sides.  Carefully lift the cabinet using a left and right rocking motion.

6.	Locate the LS-NET board and remove the screw attaching the board to the rear bezel.

7.	Lift the board out of the edge connector on the main CPU board.

8.	If unit is a standard LS140 (not an Option-L1) remove IC at location J5025 on 724-F-5 board (Long board in leftmost slot in CPU as viewed from the front of the unit.)  This can be done by inserting the blade of a flat bladed screwdriver between the IC and the socket and gently prying out the IC.

9.	Install the LS-GPIB board in the slot that the Network board was in.  Attach screw to rear bezel.

10.	Reinstall cabinet.

11.	Insert the appropriate LS-INSTALL disk (standard or L1) into the floppy disk drive and turn on power to the unit.  When the firmware installation is complete remove the disk and re-power up the unit.