Purpose: A heat sink assembly must be added to the -5 volt power supply to improve reliability.

Parts Required:

|  | Description | Part Number | Qty |
| --- | --- | --- | --- |
|  | Heat Sink | 707-241-094 | 1 |
|  | 4-40x1/2” screw | 560-440-008 | 2 |
|  | Shakeproof washer | 551-430-400 | 4 |
|  | Flat washer | 575-430-002 | 3 |
|  | Nut #4 | 581-440-001 | 2 |
|  | Thermally conductive insulator | 500-400-008 | 1 |
|  | Insulating bushing | 500-120-004 | 1 |

Description:

- Attach the long end of the bracket to the heatsink on the -5 v power supply  using a 1/2” screw and shakeproof washer through the heatsink on the supply and then through the bracket. Secure using a flat washer, shakeproof washer and nut. See figure 1.

- Apply the self adhesive thermally conductive insulator to the free end of the braket near the hole.  Attach the braket to the hole in frame using  a 1/2” screw, shakeproof washer, flat washer and insulating bushing through the braket and then through the frame.  Secure with a flat washer, shakeproof washer and nut. See figure 2.

Figure 1

Figure 2