# INSTRUCTION BOOK No. 14-7490R

A.W.A. ULTRA-HIGH FREQUENCY SIGNAL GENERATOR

TYPES 4R7490 (70-150Mc.), 5R7490 (45-100Mc.)

AND 6R7490 (20-50Mc.)

(Serial Nos. above 236)

### 47 York Street, Sydney

INSTRUCTION BOOK NO. 14-7490R

A.W.A. ULTRA-HIGH FREQUENCY SIGNAL GENERATOR

TYPES 4R7490 (70-150 M.), 5R7490 (45-100 Mc.)

And 6R7490 (25-50 Mc.)

(Serial Nos. above 286)

Amalgamated Wireless (A’sia) Limited

47 York Street

SYDNEY

10493

(i)

INDEX

SECTION    									PAGE NOS.

1		DESCRIPTION:

1.1	  Construction and Circuit Arrangements			1

1.2	  Summary of Characteristics				1

2		INSTALLATION:

2.1	Adjustment for Supply Voltage				2

2.2	Insertion of Valves					2

3		OPERATING INSTRUCTIONS:

3.1	Special Caution						3

3.2	Output Connection and Terminating Unit		3

3.3	Setting  “ZERO”					3

3.4	Setting Carrier						4

3.5	Setting Frequency					4

3.6	Controlling Output					4

4		MAINTENANCE:

4.1	Fuses							4

4.2	R-F Oscillator						4

4.3	Voltmeter						4

4.4	Typical Operating Voltages and Currents		5

5		COMPONENT SCHEDULE					6

6		EQUIPMENT SCHEDULE					9

7		DRAWINGS:							DWG. NOS.

Circuit Diagram,

U.H.F. Signal Generator, type 4R7490,

5R7490 and 6R7490 (series nos. above

286)							7490C24

Frequency Calibration Chart

Frequency Calibration Curve

Typical Circuits,

A.W.A. Terminating Unit, type R8649			8649C1

(1)

INSTRUCTION BOOK NO. 14-7490R

1. DESCRIPTION

1.1 Construction and Circuit Arrangements

These A.W.A. U.H.F. Signal Generators, types 4R7490, 5R7490 a 6R7490 employ in their r-f sections the special technique necessary the frequency band in which they operate,   The r-f oscillator  is of the Hartley type, using an acorn triode, and a low-loss coil and variable condenser form the tank circuit.   The r-f output is obtained from a mutual inductance coupling to the coil arranged in the form of a piston attenuator, capable of an absolute calibration in decibels,   An 80 Ω concentric output cable is provided,   The whole of the r-f section is enclosed in a brass box, heavily tin-plated, and all leads entering the box are well filtered, variable condenser rotor and stator are milled from solid brass, and are insulated by polystyrene and micalex respectively.

The r-f voltage developed is monitored by a diode voltmeter, and may be held              	constant by varying the plate voltage of the oscillator valve (“SET CARRIER” control).   The           	voltmeter employs two acorn diodes, one being used in conjunction with a potentiometer          	(“SET ZERO” control) to supply a “backing-off” potential.

Modulation is provided by a Colpitt's type oscillator, and  is applied to the r-f 	oscillator in such a way that the modulation depth does not vary as the plate voltage is 	adjusted.   External modulation may also be employed.

The power, supply section uses a full-wave valve rectifier and a double-section            	choke-input filter.   A mains switch, pilot light, r-f mains filter and fuses are provided.

The instrument is contained in a cast aluminium case, all components being mounted 	on two vertical panels.   The larger of these panels carries the r-f oscillator box, and a chassis 	mounting the modulator, power supply filter and voltmeter “backing-off” diode. On the front 	of this panel are mounted the frequency, output, set carrier and set  zero controls, carrier and 	modulation switches, terminals for external modulation, and the output monitoring meter, The 	smaller panel carries the r-f mains filter, mains switch, fuses, power transformer, rectifier and 	pilot light,   A partition in the case effectively screens these components from the signal 	generator proper.   The case is provided with carrying handles and with cleats on which the 	output and power cables may be wound when not in use,

1.2	Summary of Characteristics

(a)	Frequency Range

Type 4R7490:		70-150	Mc.

Type 5R7490: 		45-100	Mc.

Type CR7490: 		25-50	Mc.

(b)	Output Voltage

Maximum value:	0.1 volt. when correctly terminated

(refer 3.2)

Monitoring:	Input to attenuator held constant by manual

control and valve voltmeter.

(c)	Attenuator Range

0-100	db. below maximum output

(1)

INSTRUCTION BOOK NO. 14-7490R

SECTION 1 (Contd.)

(d)	Output Connection

80 Ω unbalanced transmission line

(e)	Modulation

Internal:	400 ~ ±10 ~ distortion less than 2%, depth                      						approximately 2%

External:	Required input - 40V R.M.S. for 30% modulation 			(approximate)

Input Impedance - approximately 20,000 to 17,000 Ω 						(resistive)

(f)	Power Supply

The instrument is normally supplied with a mains unit, series H7491, for                                                        		operation on a supply of 200-250V, 50-60~, but a special power supply unit, series                           		H50361, is available for operation from a supply of 80V, 800-3,000~.

Special units, for use with other voltages and frequencies, may also be                       		supplied, if required.

The power consumption is approximately 30 watts.

(g)		Dimensions

27” x 10 ½ ” x 11” over all projections

(h)		Weight

65 lbs.

2. INSTALLATION

2.1	Adjustment  for Supply Voltage

The Mains Unit Fuses should be placed in the correct clips to suit the mains vo1tage.   Three transformer taps are provided, namely 220, 240 and 260V.   One fuse should be in a clip marked with the voltage to be employed, and the other fuse in the clip marked common.   The fuses are accessible when the Mains Unit is removed from the case.

2.2	Insertion of Valves

The valves should be inserted in their respective sockets, which are plainly marked,            	and the top connections made where required.  The two valves in the oscillator box are accessible 	through cover plates without removing the back of the box.

(3)

INSTRUCTION BOOK NO. 14-7490R

SECTION 3   (Contd.)

3.1  SPECIAL CAUTION

CARE  MUST BE TAKEN NOT TO SUBJECT THE MAIN TUNING CONDENSER TO

UWARRANTED MECHANICAL STRAIN, AS THIS MAY RESULT IN FRACTURE  OF THE INSULATING RODS.  UNDUE PRESSURE ON THE FREQUENCY DIAL OR KNOB, ORTHE STRIKING OF EITHER WITH OPEN APPARATUS MUST BE CAREFULLY GUARDED AGAINST.

Output Connection and Terminating Unit

THE OUTPUT CABLE MUST BE TERMINATED CORRECTLY WITH ITS CHARACTERISTIC IMPEDANCE (i.e. 80 Ω RESISTANCE).   OTHERWISE LARGE FLUCTUATIONS OF VOLTAGE WILL OCCUR AS THE FREQUENCY IS VARIED, AS WELL AS ERRORS IN OUTPUT VOLTAGE  MEASUREMENT .

The terminating unit type R8649 which may be supplied, forms a convenient means of terminating the output cable, and may also be used to insert series resistors or L, T or Π pads into the output connection,   A selection of non-inductive resistors, conducting slugs and insulating slugs is supplied, and other values of resistor may be obtained,   A selection of adaptors for various types of coaxial connectors and for connection by screw-down terminals is also supplied.

The following examples will show how the Terminating Unit should be used (see Drg. No  8649C1).

(a)	Connection to High Impedance Circuit (e.g., intermediate frequency grid circuit)

Use 80 Ω shunt resistor in T-piece and adaptor as required

(b)	Connection to 80 Ω (or 100 Ω) Circuit (e.g., aerial input circuit)

Use adaptor as required, do not use T-piece.

(c)	Insertion of Attenuating Pad in 80 Ω Circuit

Use T or П-pad (one or to T-pieces respectively) and adaptor as required, e.g., a Π-pad for 40 db  attenuation at 80 Ω would comprise two 80 Ω shunt arms and a 4,000 Ω series arm.

(a)	Provision of Lo  Impedance Output

Use en inverted L-pad (one T-piece) and adaptor as required. For a 7 Ω (approx.) output (voltage x 1/10) use a 70 Ω series arm (on the generator side of the T-piece) and on 8 Ω shunt arm,

A series resistor (say 70, 80 or 100 Ω ) may be used with this L-pad (in the output arm of the T-piece) for determination of receiver input impedance.

3.3	Setting “ZERO”

With the “carrier” switch at “OFF”, the meter deflection should be zero.  It may be corrected at any time by means of the “SET ZERO” adjustment, having first set the mechanical zero of the meter with the main switch off.  If it is found that the “SET ZERO” adjustment is insufficient, one of the  diodes requires replacement (see "MAINTENANCE ).

(4)

INSTRUCTION BOOK NO. 14-7490R

SECTION 3   (Contd.)

3.4	Setting Carrier

The r-f voltage developed varies somewhat with frequency, and the meter must be at all times kept  at the “0.1V” mark by adjustment of the “"SET CARRIER” control.

3.5		The  frequency calibration is given in the Table and curve attached.

3.6	Controlling Output

The output control is calibrated to read directly in decibels below the r-f input to the attenuator, which is 0.1 when the meter reads  ran “0.1V” provided the output is correctly terminated in 80 Ω.

4.  MAINTENANCE

4.1	Fuses

Two 3-ampere fuses are mounted below the Mains Unit chassis

4.2	R-F Oscillator

Failure to obtain sufficient r-f voltage indicates that the Oscillator valve (V5, type 955) requires replacement.   Changing this valve will have a small effect on the frequency calibration, but the change is usually less than l%.

4.3	Voltmeter

If the meter deflects off scale backwards at all settings of the “SET ZERO” control, the r-f diode (V4, type EA5O) has failed and must be replaced.   This valve is mounted within the oscillator.

If the meter deflection cannot be brought down to zero by means of the “SET ZERO” control, the backing-off diode (V3, type EA5O) has failed and must be replaced,   This valve is mounted on the upper side of the chassis.

If the meter will not zero after replacing either diode, the diodes should be interchanged.   The valves used are standard commercial products, and any pair will function satisfactorily.

(5)

INSTRUCTION BOOK NO. 14-7490R

SECTION 4 (Contd.)

4.4	Typical Operating Voltages and Currents.

(a)	Mains Unit

Input Voltage	220-260V, 50-60~ (to appropriate tap)

Output plug removed from socket on generator chassis

V1 plate - earth voltages:	320V, 320V  R.M.S.

V1 heater voltage:		5V R.M.S.

L.T. Supply at output

socket				6.5V R.M.S.

(b)  	Signal Generator Proper

Connected to mains unit

by interconnecting cable,

all valves in sockets.

Valve heaters:			6.3V R.M.S.

H.T, voltage (measured

across C8):			260V (Carrier off, internal or external modulation)

Hum voltage on H.T.

(across C8) :			Less than 0.25Vpeak

V2 voltages:

Int. Mod.	Ext. Mod.

Plate- cathode.			230V	           * l0V

Screen-cathode			70V		40V

Cathode-earth			l.3V		113V

V2 currents

Int. Mod.	Ext. Mod.

Plate				4.3 mA		4.4 mA

Screen				0.7 mA		0.8 mA

Grid				0.2 mA		 Nil

Cathode			5.2 mA		5.2 mA

V5 plate supply voltages:

D.C. (junction of R12, C30 to earth)  	250V

A.F. (junction of C30, S2  to earth)	40V R.M.S.(approx.)

V5 plate  voltage (S3 to earth):		30-180 V D.C.

V5 plate current (maximum): 10 mA

*  This low voltage is due the inclusion of a series resistor in circuit in the External 	Modulation position in order to maintain the H.T. voltage unchanged.

(6)

INSTRUCTION BOOK NO. 14-7490R

5. COMPONENT SCHEDULE

U.H.F. Signal Generators, type 4R, 5R, and 6R7490

Circ.		Effective Value				A.W.A. Type Number

Ref. No.						(unless otherwise stated)

(a) 	Resistors

R1		0.25 MΩ, ¼ W				Carbon

R2		Potentiometer, 20,000 Ω		Marquis Vernier

R3		Not Used

R4		4,000 Ω, ¼ W				Carbon

R5		Not Used

R6		16,000 Ω, ¼ W				Carbon

R7		25,000 Ω, ½ W				Bifrost, Carbon

R8		75 Ω, ⅓ W				Bifrost Carbon

R9		Short circuit if V3 and V4 are type EA50: 5.75 Ω type AB, Ctg.C if D1

R10		250 Ω,					W/wound, type AB, Ctg.C

R11		0.25 MΩ, 1W				Carbon

R12		5,000 Ω					W/wound, type AB, Ctg.C

R13		Not Used

R14		….. Ω ¼ W (select on test for 30%

Modulation) (nominally 4,000 Ω)	Carbon

R15		50,000 Ω (2 x 100,000 Ω, 1W in

Parallel)				Carbon

R16		Potentiometer, 25,000 Ω		Carbon

R17		Not Used

R18		Not Used

R19		75 Ω, ⅓W				Bifrost, Carbon

R20		Not Used

R21		Not Used

R22		5,000Ω					W/wound, type AB, Ctg.C

(b)	Condensers

C1		1,000 µµF, mica			Simplex P/T

C2		1,000 µµF, mica			Simplex P/T

(7)

INSTRUCTION BOOK NO. 14-7490R

SECTION 5 (Contd.)

Circ.		Effective Value				A.W.A. Type Number

Ref. No.						(unless otherwise stated)

C3		1,000 µµF, mica			Simplex P/T

C4		1,000 µµF, mica			Simplex P/T

C5		1,000 µµF, mica			Simplex P/T

C6		Not Used

C7		5 µF, 350 V.W. Paper			3U3551

C8		5 µF, 350 V.W. Paper			3U3551

C9		.05 µF, 350 V.W. Paper			S7080

C10		.1 µF, 350 V.W. Paper			S7080

C11		.25 µF, 350 V.W. Paper			S7080

C12		1,000 µµF, mica			Simplex P/T

C13		Not Used

C14		1,000 µµF, mica			Simplex P/T

C15		Not Used

C16		1,000 µµF, mica			Simplex P/T

C17		1,000 µµF, mica			Simplex P/T

C18		1,000 µµF, mica			Simplex P/T

C19		Not Used

C20		1,000 µµF, mica			Simplex P/T

C21		Not Used

C22		1,000 µµF, mica			Simplex P/T

C23		1,000 µµF, mica

C24		Not Used

See note in each

individual book

C25		1,000 µµF, mica			Simplex P/T

C26)		10 µµF, mica (for type 4R7490)		Simplex D/S

)		100 µµF, mica (for type 5R7490)	Simplex D/S

)		500 µµF, mica (for type 6R7490)	Simplex D/S

)

C27)		5 µµF, mica (for type 4R7490)		Simplex D/S

)		10 µµF, mica (for type 5R7490)		Simplex D/S

)		50 µµF, mica (for type 6R7490)		Simplex D/S

)

C28)		Variable, 4-30 µµF ( for type 4R4790)

)		Variable 5-48 µµF ( for type 5, 6r4790)

)

C29			0.25 µF, 350V.W., paper		S7080

(8)

INSTRUCTION BOOK NO. 14-7490R

SECTION 5 (Contd.)

Circ.		Effective Value				A.W.A. Type Number

Ref. No.						(unless otherwise stated)

C30

C31		1,000 µµF, mica			Simplex P/T

C32

C33			1,000 µµF, mica			Simplex P/T

C34			1,000 µµF, mica			Simplex P/T

(c)	Radio Frequency Chokes

RFC1,2,	R.F. chokes, Dwg. No.

4-6, 8-11	8483T51-1 (resis. 1Ω)			4R7490

RFC1,2,	R.F. chokes, Dwg. No.

4-6, 8-11	8483T51 (resis. 1Ω)			5R,6R7490

RFC3 & 7	Not used

(d)	Power Transformers and Chokes

T1	Power transformer, (resis-

tances: Primary 25Ω, H.T.

800Ω C.T., L.T. 0Ω				1TJ7447

L1	Filter choke (resistance 500Ω)			1TX7446

L2	Filter choke (resistance 500Ω)			1TX7446

(e)	Audio-Frequency Chokes

L3	Feed choke (resistance 500Ω)			1TX7446

L4	Plate-grid choke (resistance 250Ω)		5TA4793

(f)	Valves

V1	Rectifier					5Y3G

V2		A.F. Oscillator					6V6G

V3		Backing-off diode, type EA50 or

type D1 (N.B.: see R9)

V4		Metering diode, type EA50 or

type D1 (N.B.: see R9)

V5		R.F. oscillator					955

(g)	Switches

S1		Mains switch					S3510

S2		Modulation switch (Dwg. No. 7490D7)		Oak type 28

S3		Carrier switch (Dwg. No. 7490D3)		Oak type 25

(9)

INSTRUCTION BOOK NO. 14-7490R

SECTION 5 (Contd.)

Circ.		Effective Value				A.W.A. Type Number

Ref. No.						(unless otherwise stated)

(h)	Meter

M1		0-50 µA (resistance

Approx. 1,000 Ω)

(j)	Miscellaneous

F1		Fuse 3A				Glass cartridge type

F2		Fuse 3A				Glass cartridge type

PL1		pilot lamp 6.3V				Mazda type 46

Oscillator Box						(1A8483 (for type 4R7490)

(2A8483 (for type 5R7490)

(3A8483 (for type 6R7490)

Attenuator						1R7493

Output Connector					(7490U261 (for type 6R7490)

(HMV type D4020C (equiv.

Pye elbow (for type 4,

(5R,7490)

Mains Unit						(7H7491 (for type 4R7490)

(8H7491 (for type 5, 6R7490)

6. EQUIPMENT SCHEDULE

Item 1.		One A.W.A. U.H.F. Signal Generator, type 4R7490, 5R7490 or 6R7490.

Item 2.		One A.W.A. Mains Unit, type 7H7491 (for type 4R7490) or 8H7491 (for types

5R7490 and 6R7490) mounted with item 1.

Item 3.		Valves for the operation of items 1 and 2 as follows :-

One type 5Y3G

One type 6V6G

One type 955

Two type EA50 or

Two type MAZDA type D1

Item 4.			Instruction Book No. 14-7490R

Note:	This list is a guide to the items which comprise a normal equipment, but may                        			be varied to suit the requirement of each installation.

A.W.A. U.H.F. SIGNAL GENERATOR

TYPE 6R7490 SERIAL NO. 524

NAME PLATE

………………………..

……………………….

………………………….

Frequency Calibration

| Frequency (Mc.) | Dial Setting |
| --- | --- |
| 22 | 3.85 |
| 24 | 13.05 |
| 25 | 17.2 |
| 26 | 20.95 |
| 28 | 27.75 |
| 30 | 33.8 |
| 32 | 39.45 |
| 34 | 44.75 |
| 35 | 47.75 |
| 36 | 50.0 |
| 38 | 55.05 |
| 40 | 60.25 |
| 42 | 65.4 |
| 44 | 71.05 |
| 45 | 73.7 |
| 46 | 76.55 |
| 48 | 83.4 |
| 50 | 91.75 |