1.

INSTRUCTION BOOK N0. 3-7490R

A.W.A. ULTRA-HIGH FREQUENCY SIGNAL GENERATORS

TYPES 8R7490 & 9R7490

1. DESCRIPTION

1.1 Construction and Circuit Arrangements:

This Instruction Book describes the A.W.A  U.H.F. Signal Generators bearing the following name plates.  The corresponding A.W.A. Type numbers are indicated since they do not appear on the nameplate.

Nameplates and Type Numbers

| U.H.F. SIGNAL GENERATOR |  | MAINS UNIT |  |
| --- | --- | --- | --- |
| Nameplate | A.W.A. Type | Nameplate | A.W.A. Type |
| Signal Generator 200 Mc. Mk Serial No. D↑D | 8R7490 | Mains Unit for Signal Generator 200 Mc. Mk Serial No. D↑D | 1H7491 |
| R.A.A.F. Test Set Type A6 Ident. No. Y10S/60,006 Serial No. | 9R7490 | R.A.A.F. Mains Unit A6A Ident. No. Y10S/60,013 Serial No. | 2H7491 |

The A.W.A. U.H.F. Signal Generator, types 8R7490 and 9R7490 employs in its r-f sections the special technique necessitated  by the frequency band in which they operate.   The r-f oscillator  is of the Hartley type, using an acorn triode, and a short circuited transmission line forms the tank circuit.   The line is loaded by a variable capacity to provide the frequency adjustment.  The r-f output is obtained from a mutual inductance coupling to the line arranged  in the form of a piston attenuator, capable of an absolute calibration in decibels.   An 80 Ω concentric output cable is provided.   The whole of the r-f section is enclosed in a brass box, heavily silver-plated, and all leads entering the box are well filtered. The  variable condenser rotor and stator are milled from solid brass, and are insulated by polystyrene and micalex respectively.

The r-f voltage developed is monitored by a diode voltmeter, and may be held              	constant by varying the plate voltage of the oscillator valve (“SET CARRIER” control).   The           	voltmeter employs two acorn diodes, one being used in conjunction with a potentiometer          	(“SET ZERO” control) to supply a “backing-off” potential.

Modulation is provided by a Colpitt's type oscillator, and  is applied to the r-f 	oscillator in such a way that the modulation depth does not vary as the plate voltage is 	adjusted.   External modulation may also be employed.

The power, supply section uses a full-wave valve rectifier and a double-section            	choke-input filter.   A mains switch, pilot light, r-f mains filter and fuses are provided.

2.

INSTRUCTION BOOK N0. 3-7490R

A.W.A. ULTRA-HIGH FREQUENCY SIGNAL GENERATORS

TYPES 8R7490 & 9R7490

SECTION 1 (Contd.)

The instrument is contained in a cast aluminium case, all components being mounted 	on two vertical panels.   The larger of these panels carries the r-f oscillator box, and a chassis 	mounting the modulator, power supply filter and voltmeter “backing-off” diode and spare                      	valves. On the front of this panel are mounted the frequency, output, set carrier and set             	zero controls, carrier and modulation switches, terminals for external modulation, and                  	the output monitoring meter, The smaller panel carries the r-f mains filter, mains                 	switch, fuses, power transformer, rectifier and 	pilot light,   A partition in the case effectively 	screens these components from the signal generator proper.   The case is provided with 	carrying handles and with cleats on which the output and power cables may be wound when 	not in use.

1.2	Summary of Characteristics

(a)		Frequency Range

140-300 Mc.

(b)	Output Voltage

Maximum value:  About 0.4 volt.

Monitoring:  Input to attenuator held constant by manual control and valve voltmeter.

(c)	Attenuator Range

0-100	db. below maximum output

(d)	Output Connection

80 Ω unbalanced unterminated transmission line

(e)	Modulation

Internal: 400 c/s ±10 c/s. distortion less than 2%, depth approximately 2%

External:	Required input - 40V R.M.S. for 30% modulation (approximate)

Input Impedance - approximately 20,000 to 17,000 Ω (resistive)

(f)	Power Supply

220-260V, 50-60 c/s, approximately 30 watts.

3.

INSTRUCTION BOOK N0. 3-7490R

A.W.A. ULTRA-HIGH FREQUENCY SIGNAL GENERATORS

TYPES 8R7490 & 9R7490

SECTION 1 (Contd.)

(g)		Dimensions

27” x 10 ½ ” x 11” over all projections

(h)		Weight

65 lbs.

4.

INSTRUCTION BOOK N0. 3-7490R

A.W.A. ULTRA-HIGH FREQUENCY SIGNAL GENERATORS

TYPES 8R7490 & 9R7490

2. INSTALLATION

2.1	Adjustment for Supply Voltage

The Mains Unit fuses should be placed in the correct clips to suit the mains voltage.  Three transformer taps are provided, namely 220, 240 and 260V.  One fuse should be in a clip marked with the voltage to be employed, and the other fuse in the clip marked common.  The fuses are accessible when the Mains Unit is removed from the case.

2.2 	Insertion of Valves:

The valves should be inserted in their respective sockets, which are plainly marked, and top connections made where required.  The two valves in the oscillator box are accessible through cover plates without removing the back of the box.

5.

INSTRUCTION BOOK N0. 3-7490R

A.W.A. ULTRA-HIGH FREQUENCY SIGNAL GENERATORS

TYPES 8R7490 & 9R7490

3.1  SPECIAL CAUTION

CARE  MUST BE TAKEN NOT TO SUBJECT THE MAINTUNING CONDENSER TO

UWARRANTED MECHANICAL STRAIN, AS THIS MAY RESULT IN FRACTURE  OF THE INSULATING RODS.  UNDUE PRESSURE ON THE FREQUENCY DIAL OR KNOB, ORTHE STRIKING OF EITHER WITH OPEN APPARATUS MUST BE CAREFULLY GUARDED AGAINST.

Output Connection

The output lines must be terminated correctly (i.e., with 80 Ω resistance), otherwise                   		large fluctuations of voltage will occur as the frequency is varied.

3.3	Setting “ZERO”

With the “carrier” switch at “OFF”, the meter deflection should be zero.  It may be                               	corrected at any time by means of the “SET ZERO” adjustment, having first set the            	mechanical zero of the meter with the main switch off.  If it is found that the “SET ZERO”                   	adjustment is insufficient, one of the  diodes requires replacement (see "MAINTENANCE ).

3.4	Setting Carrier

The r-f voltage developed varies somewhat with frequency, and the meter must be at all times kept  at the “40µA – 0.25V” mark by adjustment of the “"SET CARRIER” control.

3.5		The  frequency calibration is given in the Table and curve attached.

3.6	Controlling Output

The output control is calibrated to read directly in decibels below the r-f input to the attenuator.

6.

INSTRUCTION BOOK N0. 3-7490R

A.W.A. ULTRA-HIGH FREQUENCY SIGNAL GENERATORS

TYPES 8R7490 & 9R7490

4.  MAINTENANCE

4.1	Fuses

Two 3-ampere fuses are mounted below the Mains Unit chassis

4.2	R-F Oscillator

Failure to obtain sufficient r-f voltage indicates that the Oscillator valve (V5, type 955) requires replacement.   Changing this valve will have a small effect on the frequency calibration, but the change is usually less than l%.

4.3	Voltmeter

If the meter deflects off scale backwards at all settings of the “SET ZERO” control, the r-f diode (V4, type EA5O) has failed and must be replaced.   This valve is mounted within the oscillator.

If the meter deflection cannot be brought down to zero by means of the “SET ZERO” control, the backing-off diode (V3, type EA5O) has failed and must be replaced,   This valve is mounted on the upper side of the chassis.

If the meter will not zero after replacing either diode, the diodes should be interchanged.   The valves used are standard commercial products, and any pair will function satisfactorily.

4.4	Typical Operating Voltages and Currents:

(a)	Mains Unit

Input Voltage:			220-260V, 50-60 cycles

(to appropriate tap)

Output plug removed from

Socket.

V1 plate - earth voltages:	340V, 340V  R.M.S.

V1 heater voltage:		5V R.M.S.

L.T. Supply at output

socket				6.5V R.M.S.

(b)  	Signal Generator Proper.

Connected to mains unit by interconnecting cable,

all valves in sockets.

Valve heaters:			6.3V R.M.S.

H.T, voltage (measured 	across C8): 260V (Carrier off, internal or external             		modulation)

Hum voltage on H.T.

(across C8) :			Less than 0.25Vpeak

7.

INSTRUCTION BOOK N0. 3-7490R

A.W.A. ULTRA-HIGH FREQUENCY SIGNAL GENERATORS

TYPES 8R7490 & 9R7490

V2 voltages:

Int. Mod.	Ext. Mod.

Plate- cathode.			260V	           * l0V

Screen-cathode			70V		40V

Cathode-earth			l.3V		1.3V

V2 currents:

Int. Mod.	Ext. Mod.

Plate				4.3 mA		4.4 mA

Screen				0.7 mA		0.8 mA

Grid				0.2 mA		 Nil

Cathode			5.2 mA		5.2 mA

V5 plate supply voltages:

D.C. (junction of R12, C30 to earth)  	250V

A.F. (junction of C30, S2  to earth)	40V R.M.S.(approx.)

V5 plate  voltage (S3 to earth):		0-100 V D.C.

V5 plate current (maximum): 5 mA

*  This low voltage is due the inclusion of a series resistor in circuit in the External 	Modulation position in order to maintain the H.T. voltage unchanged.

8.

INSTRUCTION BOOK N0. 3-7490R

A.W.A. ULTRA-HIGH FREQUENCY SIGNAL GENERATORS

TYPES 8R7490 & 9R7490

IMPORTANT !

HOW TO ORDER REPLACEMENT PARTS

The following systems of ordering spare or replacement parts should be adhered to in order to avoid unnecessary correspondence and expedite delivery.

1.	To order a major part which carries its own type number, quote this type number;                       	quote also serial number, resistance value or capacity value where shown.

Examples:-	Transformer 1TK5510 (Serial No. 117)

Resistance A.W.A. 1S4282  500Ω

2.	To order a small electrical part which carries no type number, quote:-

(i)	Type number and serial number of section (not whole equipment) where part                                		is used.

(ii)	Circuit reference number.

(iii)	Circuit diagram number and sub-number, the latter being found in the right-                           		hand section of the last entry in the “changes” column.

Examples:-	Type PS368, serial, No. 18, C50A Dwg. No. 8368C1,sub-number 2.

3.	To order a part which carries no type number and has no circuit reference number,                       	quote:-

(i)	Type number and serial number of section (not whole equipment) where part                                		is used.

(ii)	 Circuit reference number and circuit dwg. Number of component associated                       		with the wanted part.

(iii)	Brief description of part.

Examples:- Type J8279, Serial No. 41.  Insulated coupling C23B, Dwg. No. 8279D1.

5. COMPONENT SCHEDULE

Circ.		Effective Value				A.W.A. Type Number

Ref. No.						(unless otherwise stated)

(a) 	Resistors

R1		0.25 MΩ, ¼ W

R2		Potentiometer, 20,000 Ω		Marquis Vernier

R3		Not Used

R4		4,000 Ω, ¼ W

R5		Not Used

R6		16,000 Ω, ¼ W

9.

INSTRUCTION BOOK N0. 3-7490R

A.W.A. ULTRA-HIGH FREQUENCY SIGNAL GENERATORS

TYPES 8R7490 & 9R7490

Section 5 (Contd.)

Circ.		Effective Value				A.W.A. Type Number

Ref. No.						(unless otherwise stated)

R7		25,000 Ω, ½ W				Bifrost, Carbon

R8		75 Ω, ⅓ W				Bifrost Carbon

R9		Short circuit if V3 and V4 are Mullard

type EA50: 5.75 Ω if Mazda D1

R10		250 Ω,	3Watt

R11		0.25 MΩ, 1W

R12		10,000 Ω 5Watt

R13		Potentiometer, 200 Ω			AWA S5711

R14		….. Ω ¼ W (select on test for 30%

Modulation) (nominally 4,000 Ω)

R15		50,000 Ω 2W

R16		Potentiometer, 25,000 Ω

R17		Not Used

R18		Not Used

R19		30 Ω, ⅓W				Bifrost

R20		30 Ω, ⅓W				Bifrost

(b)	Condensers

C1		0.002 µF, mica

C2		0.002 µF, mica

C3		0.002 µF, mica

C4		0.002 µF, mica

C5		0.001 µF, mica

C6		0.001 µF, mica

C7		5 µF 350 V.W.				AWA 3U3551

C8		5 µF 350 V.W.				AWA 3U3551

C9		0.05 µF 350 V.W.			AWA S7080

C10		0.1 µF 350 V.W.			AWA S7080

C11		0.25 µF 350 V.W.			AWA S7080

C12		0.001 µF, mica

C13		Not Used

C14		0.001 µF, mica

10.

INSTRUCTION BOOK N0. 3-7490R

A.W.A. ULTRA-HIGH FREQUENCY SIGNAL GENERATORS

TYPES 8R7490 & 9R7490

Section 5 (Contd.)

Circ.		Effective Value				A.W.A. Type Number

Ref. No.						(unless otherwise stated)

C15		115 µµF Silvered Mica

C16		115 µµF Silvered Mica

C17		115 µµF Silvered Mica

C18		115 µµF Silvered Mica

C19		115 µµF Silvered Mica

C20		200 µµF Silvered Mica

C21		115 µµF Silvered Mica

C22		200 µµF Silvered Mica

C23		115 µµF Silvered Mica

C24		115 µµF Silvered Mica

C24		115 µµF Silvered Mica

C25		115 µµF Silvered Mica

C26	 	10 µµF Silvered Mica

C27		5 µµF Silvered Mica

C28		Variable, 5 to 32 115 µµF 			AWA 1U7492

C29			0.25 µF, 350V.W., paper			AWA S7080

C30			0.5 µF, 350V.W., paper				AWA S7080

C31		115 µµF Silvered Mica

C32			0.001 115µF Mica

(c)	RFC1-5		R.F. Chokes, Drawing 7490T94 Section 5 (Contd.)

RFC6		R.F. Choke Drawing 7490T94-1 Section 5 (Contd.)

(d)	Power Transformers and Chokes

T1	Power transformer, (resis-

tances: Primary 25Ω, H.T.

800Ω C.T., L.T. 0Ω				1TJ7447

L1	Filter choke (resistance 500Ω)			1TX7446

L2	Filter choke (resistance 500Ω)			1TX7446

(e)	Audio-Frequency Chokes

L3	Feed choke (resistance 500Ω)			1TX7446

11.

INSTRUCTION BOOK N0. 3-7490R

A.W.A. ULTRA-HIGH FREQUENCY SIGNAL GENERATORS

TYPES 8R7490 & 9R7490

Section 5 (Contd.)

Circ.		Effective Value				A.W.A. Type Number

Ref. No.						(unless otherwise stated)

L4	Plate-grid choke (resistance 250Ω)		5TA4793

(f)	Valves

V1	Rectifier					Radiotron 5Y3G

V2		A.F. Oscillator					Radiotron 6V6G

V3		Backing-off diode, type EA50 or

type D1 (N.B.: see R9)

V4		Metering diode, type EA50 or

type D1 (N.B.: see R9)

V5		R.F. oscillator					Radiotron 955

(g)	Switches

S1		Mains switch					S3510

S2		Modulation switch (Dwg. No. 7490D7)		Oak type 28

S3		Carrier switch (Dwg. No. 7490D3)		Oak type 25

(h)	Meter

M1		0-50 µA (resistance

Approx. 1,000 Ω)

(j)	Miscellaneous

F1		Fuse 3A				Glass cartridge type

F2		Fuse 3A				Glass cartridge type

PL1		pilot lamp 6.3V				Mazda type 46

Oscillator Box						AWA A8483

Attenuator						1R7493

Output Connector:-	U.H.F. Signal Generator Type 8R7490: Amphenol Type 93M

U.H.F. Signal Generator Type 8R7490: Pye elbow

Mains Unit		U.H.F. Signal Generator Type 8R7490: 7H7491

U.H.F. Signal Generator Type 9R7490: 2H7491

Interconnection Cable – Dwg. 7490U131

12.

INSTRUCTION BOOK N0. 3-7490R

A.W.A. ULTRA-HIGH FREQUENCY SIGNAL GENERATORS

TYPES 8R7490 & 9R7490

6. EQUIPMENT SCHEDULE

Item 1.		One A.W.A. U.H.F. Signal Generator, type 8R7490 or 9R7490.

Item 2.		One A.W.A. Mains Unit, type 17H7491 (for type 8R7490) or 2H7491 (for type

9R7490 mounted with item 1.

Item 3.		Valves for the operation of items 1 and 2 as follows :-

One Radiotron type 5Y3G

One Radiotron type 6V6G

One Radiotron type 955

Two Mullard type EA50 or

Two Mazda type D1

Item 4.			Instruction Book No. 3-7490R

Note:	This list is a guide to the items which comprise a normal equipment, but may                        			be varied to suit the requirement of each installation.

W.A. U.H.F. SIGNAL GENERATOR

TYPE 8R7490 SERIAL NO. 24

Name Plate

………………………………………..

……………..200Mc………………….

…………..……D↑D……………….…

.…………….Ser No. 24 ………………

Frequency Calibration

| Frequency (Mc.) | Dial Setting |
| --- | --- |
| 140 | 4.8 |
| 150 | 11.35 |
| 260 | 17.1 |
| 170 | 22.25 |
| 180 | 26.95 |
| 190 | 31.3 |
| 200 | 35.35 |
| 210 | 39.4 |
| 220 | 43.35 |
| 230 | 47.25 |
| 240 | 51.1 |
| 250 | 55.2 |
| 260 | 59.5 |
| 270 | 64.1 |
| 280 | 69.65 |
| 290 | 76.45 |
| 300 | 86.1 |