1.

DESCRIPTION

Application.

A.W.A. Voltohmyst type A56010 is a versatile instrument which has application in the measurement of positive and negative D.C. voltages, D.C. current, resistance and A.C. voltages.

The instrument features a circuit which measures directly the peak-to-peak values of sinusoidal and complex waveforms, the scales being calibrated directly in peak-to-peak values and also in R.M.S. for sinusoidal waveforms.

The instrument has a rated input for non-symmetrical wave-forms of 2,100V, but sinusoidal and symmetrical waveforms of up to 4,200V. can be measured.  (Refer to figure 1 below for waveforms measurable with the instrument.)

Fig.1 Waveforms Measurable with the Instrument

When measuring voltages, with a source impedance of approximately 100 ohms, the A.C. ranges up to and including 0-500V. are frequency compensated to provide a flat characteristic from 30 cycles to 2.5 Mc.  When used in conjunction with a crystal diode probe such as A.W.A. type 2R56020, the input capacitance is reduced to a low value and the A.C. voltage ranges are made accurate to within 10% from 50 kc. to 250 Mc.   (Refer to subsection 1.4.3 for frequency response figures.)

When using an associated D.C. probe, on all D.C. ranges, the instrument has an input resistance of 11 M.ohms, thus offering considerable advantages over conventional test equipment in a number of applications such as those listed below:-

2.

(i) 	Oscillator Grid-bias

The negative D.C. voltage developed at the grid of an oscillator is proportional to the amplitude of oscillation.

The instrument, when connected directly to the grid, measures this bias, with negligible effect on the conditions of oscillation.

When checking multi-band receivers, comparative readings should be taken, rotating the oscillator tuning control through its entire range during each check.

(ii)	A.V.C. Voltage

The instrument has negligible effect on A.V.C. measurements, allowing them to be taken directly at the grids of the stages to which the A.V.C. is applied, or on the A.V.C. line or at the output of the diode load.

(iii)	Output Meter

In the alignment of TV, AM or FM receivers the instrument can be employed as an output meter.   The centre-zero feature is particularly useful in aligning the discriminators of the latter receivers (refer to sub-section 2.4 of this book).

(iv)	Detection of Gassy Valves and Leaky Coupling Capacitors

The presence of an abnormally low value of negative bias or of a definite positive bias at the grid of a stage may indicate a gassy valve or leaky coupling capacitor;  this fault is readily located with the aid of the Voltohmyst.

(v)	Insulation Resistance Measurement

High insulation resistances such as those of capacitors, transformers etc. can be measured by applying an external voltage and connecting the instrument as indicated in fig. 4.(Refer also to sub-section 2.6.)

Additional applications of the instrument are:-

(i)	Measurement of dbm

By using the instrument in conjunction with the dbm conversion chart 		(refer 	to fig. 5), it is possible to determine the dbm. values corresponding to 		R.M.S. values 	measured across a 600-ohm resistive circuit.

3.

(ii)	Peak -to-peak pulse values

The peak-to-peak values of pulse waveforms may be measured, using the A.C.                                              		peak-to-peak scale of the instrument.   A long time-constant is incorporated in                                      		the circuits to allow for measurement of signals having low repetition rates.

Further information on facilities and electrical characteristics is given in sub-section

1.2	Mechanical construction

The instrument is housed in a compact metal case, having meter, four controls and 	three terminals mounted symmetrically on the front.   An engraved front panel is 	fitted, and a carrying handle provided at the top of the case.

Dimensions are as follows:-

Height:-	    10"

Width:-	6 ¼"

Depth:	6 ¼"

Weight:	 8 ½ lbs

1.3	Design  Characteristics and Circuit Description

The circuit is designed to give a high degree of accuracy in all applications, making 	the instrument especially adaptable for work in television and industrial equipment.

Maximum peak-to-peak input rating for non-symmetrical complex waveforms is 	2,100V. and for sinusoidal or symmetrical waveforms up to 4,200V.

In the case of sine waves, the instrument is calibrated directly for R.M.S. and peak-	to-peak values (a feature made practicable by the constant ratio of 2.83:1 of these 	values). The peak-to-peak-reading feature has particular application in servicing 	television receivers when taking measurements of wave-forms as indicated in Fig.

It should be noted that in this instrument, as in all voltage measuring devices, certain 	precautions are necessary in interpreting the readings, and inaccuracies may be due to 	external conditions such as high impedance of the circuit being measured, high-	frequency voltage source, or if the voltage waveform consists of short pulses with a 	low repetition rate.   The  error introduced increases with the ratio of the source

4

impedance to the impedance of the instrument, and varies with frequency.   Refer to para. 1.4.3 (c). The error in short-pulse measurement when the repetition rate is low is due to the time constants involving the capacitors in the peak-to-peak rectifier in the instrument.   These capacitors do not have time to become fully charged during the pulse, and in the relatively long period between pulses, are unable to retain the charge they have received.   The range or conditions allowing satisfactory measurement of essentially rectangular pulses derived from a 50 ohms source indicated in figure 2 below.  For sources of higher impedance error will be greater than 2%.

Fig.2	Pulse Response Capability

The instrument is frequency-compensated for A.C. voltage ranges up to and including 0-500V.   On the higher range the response varies with frequency.

All resistors in the divider networks associated with the voltage and ohms applications are high stability carbon types, having a tolerance of 1%, and in conjunction with the D.C. amplifier, ensure reliable service for a considerable period.

The fundamental element of the circuit is a D.C. bridge involving  type l2AU7 valve (V3) and its associated network. - Refer to the simplified schematic in fig. 3 below.

5.

Fig.  3

Assuming the instrument is switched on, power is supplied from the power transformer (T1), rectified by one half of a type 6AL5 valve (V2) and applied to the bridge (V3).

Providing no readings are being taken, and the circuit has been previously balanced by adjustment to the balance resistor (R25), the cathode currents of the two triode sections of the valve (V3) are equal, and the microammeter connected between the anodes reads zero.

When the facility selector (s2) is set in any of the voltage or ohms positions, and a reading is taken, a voltage is applied via range selectors, etc., to the grid of the left-hand triode (as drawn on the simplified schematic) section of V3, this altering the current through this section and in turn changing the potential across the common cathode resistor (R25).   The change in the latter voltage then acts as a reverse change in the bias on the second section, and the meter responds to the resulting changes in potential at the anodes due to the voltage drops across anode resistors R20 and R22, respectively.

When used as a milliammeter, the facility selector (S2) disconnects the meter from the electronic circuit, and connects the "+" terminal of the meter to the "+" mA-OHMS input terminal, and the "-" terminal to the  "-" mA. EARTH input, which is then disconnected from

6

earth.   Operation is then as for a conventional milliammeter, but it should be noted that all electronic protection for the meter is now out of circuit, and care must be taken not to overload.

1.4	Summary of Controls, Facilities and Electrical Characteristics

1.4.1	Facilities

Ranges

"+" D.C. Volts	)		0 – 1.5V. (On separate LO Scale for A.C.)

)

"-" D.C. Volts	)		0 – 5V.

)		0 –15V.

A.C. (R.M.S. volts	)		0 – 50V.

)		0 – 150V.

0 – 500V.

0 – 1,500V.

A.C. Volts, Peak-to-Peak;  	0 – 4.2V	(On separate LO scale)		(read simultaneously with 	0 – 14V.

R.M.S. values for sine waves) 	0 – 42v.

on separate scale.	0 – 140V.

0 – 420V.

0 – 1,4000V.

0 – 4,200V.

Current Measurement -	mA. Ranges

0 –  1.5 mA..

0 – 5 mA.

0 – 15 mA.

0 – 50 mA.

0 – 150 mA.

0 – 500mA.

0 – 1,500mA.

.

Ohms Measurement:-	Ranges

(meter calibrated	Rxl

0-1,000 ohms)	Rx 10

Rx 100

Rx 1000

Rx 10,000

Rx 100,000

Rx 1 M,ohm

7.

1.42	Controls and Input Connections

(i)	Facility Selector:	Selects desired facility, i.e., OHMS, A.C., etc.

(ii)	Range Selector:	Selects desired range, i.e., l.5V., 5V., etc.

(iii)	ZERO ADJ.		Positions the meter needle to "0" or "-0+" (the                                                                						latter when the facility selector is set to "+"                                                                						VOLTS position, and the instrument is used for                                   						centre-zero indication).

(iv)	OHMS ADJ.		Positions the meter needle at the extreme right-                                     						hand mark of the ohms scale when the facility                					   	selector is set to OHMS position.

(v)	VOLTS, A.C.-D.C.	Positions the meter needle at the extreme right-                                  						hand mark of the ohms scale when the facility                                              						selector is set to OHMS position.

(vi)	"+" mA. – OHMS	Positive input terminal for current readings.                           						Also associated with ohms facility.

(vii)	-mA. EARTH.		"-" terminal for current readings; earthed for                                            							other facilities. Second terminal for ohms                                   							measurement.

1.4.3	Electrical Characteristics

(a)	D.C. Voltmeter

Sensitivity:		Input resistance 11 megohms on all ranges

(on 1.5V. range)

Accuracy:		 3% of full scale deflection

(all ranges)		(F.S.D.)

(b)	Milliammeter

Accuracy on 		 3% of F.S.D.

all ranges

8.

(c)	A.C. Voltmeter

Input resistance and

Capacitance (with direct

Probe and cable):

1.5V., 5v., 50V., and		0.83 M.ohm, in parallel with 90 uuf.

150V.:-

500V. range :-			1.3 M.ohm in parallel with 85 uuf

1,500V. range:-		1.5 M.ohm in parallel with 85 uuf.

Frequency Response

(using direct probe and cable):-

Source Impedance		Freq. Range			Response

100 ohm			30 cycles to2.5 Mc.		+ 1db at 2.5 Mc.

1000 ohm			30 cycles to 500 kc.		+ 1db at 500 kc.

5000 ohm			30 cycles to 270 kc.		- 1db at 270 kc.

10,000 ohm			30 cycles to 80 kc.		- 1 db at 80 kc.

Frequency Response

Using crystal			Within   1db from 50 kc. To 250 Mc.

Diode probe type

Accuracy:-

1.5V., 15V., 50V.,

150V., and

500V.				 5% off F.S.D at 50 cycles

5V.				+ 0 – 10% of F.S.D. at 50 cycles.

9.

(d)	Maximum Input Voltages

Pure D.C. (i.e. without A.C. components):~

Using D.C. probe IR56020			1,500V.

Using High Voltage probe 2R56020		30,000V

A.C. Voltages  (i.e., without D.C. components):-

R.M.S. sinusoidal				1, 500V.

Peak-to-peak sinusoidal			4,200V.

Peak-to-peak

Complex waveforms				2,l00V.

Sum of D.C. and A.C. peak-to-peak voltage inputs:-

When measuring

A.C. component:-				2,100V.

D.C. component:-				1,500V

(e)	Metter Sensitivity

200 UA. for F.S.D.

(f )	Mains Supply and Power Consumption

Mains Input:-				240V., 50-60cycle,                    								single-phase.

Power Consumption:- 			5.5W. approximately.

1.5	Probes and Cables used with Voltohmyst A56010

Refer to fig. 8 (Drg. 56010D7) at the rear of this book.

1.5.1	Cables and Probes Supped with the Instrument

The instrument is normally supplied fitted with three probes (ohms probe,  		direct probe type R56020, and D.C. probe type 1R56020), together with an earth 		cable which is fitted with an alligator clip.

10.

The ohms probe and direct probe are each connected to individual cables, and are thus used alone for the measurement of resistance and A.C. voltages respectively; the circuit is completed in each case via the earth cable and clip.

The D.C. probe, however, consists of an extension piece which fits over the direct probe and clamps into place on the latter.  The whole is then used for measurement of D.C. voltages.  An alligator clip to fit onto the end of the direct probe is supplied.

1.5.2 	Probes Not Normally Supplied with the Instrument ( available on separate order)

(a)	Crystal Diode Probe 2R56020

This probe consists of an extension piece similar to the D.C. probe, with a clip attached to the side, and contains a small diode rectifier element, coupling capacitor and isolating resistor.   (Refer to the insert on the circuit diagram Drg. 5601002).

The probe has application in the measurement of R.M.S. values of sinusoidal waveforms in the range 50 kc. to 250 Mc.   The R.M.S. values are read on the D.C. scale of the Voltohmyst, and the peak values may then be obtained by multiplying by 1.414.

The probe may be used in the presence of D.C. voltages as high as 250V.

The specification is as follows:-

Frequency Response :- ± 1 db from 50 kc. to 250 Mc.

Input Capacitance :- 1.75 µµF. (approx.)

Input Resistance

Frequency		Resistance

50 kc.			370,000 ohms

100 kc.			360,000 ohms

1 Mc.			312,000 ohms

10 Mc.			283,000  ohms

100. Mc.		90,000	ohms

220 Mc.		6,000	ohms

11.

Maximum Input Voltages

R.F. Voltages		( 20V. R.M.S.

(

( 28V. peak.

D.C. Voltage		250V~.

Vo1tohmyst  Meter Indication

Sinusoidal Waveforms:		R.M.S. values read

off the DC – R.M.S.

scales.

Complex Waveforms :-		Peak values of

approximately 1.414

of indicated R.M.S. 											value for the positive 											portion of the cycle

Method of Use

(i)   	With the  probe (type 2R56020 over the tip of the direct probe.   Set the facility 		selector on the Voltohmyst to "-" VOLTS position and the range selector to an 		appropriate scale.

(ji)	Connect the clip to the earth side of the voltage source and apply the probe to the other side.

High Vo1tage Probe type 3R56020

This  probe consists of  cable and hollow-moulded piece containing a dropping resistor capable of extending the D.C range of the instrument to 0-30,000V., and is indispensable in television and heavy-duty industrial work.

The maximum input voltage usable with the probe is 30,000V., and since the probe effectively provides the D.C. ranges of the instrument with a multiplying factor of 100, the range selector may be set as required and  a direct reading obtained.  Note that the total input resistance of the Voltohmyst when using the high voltage probe in the vicinity of 1100 M. ohms, providing a valuable facility when working in high impedance circuits with poor regulation.

12.

Method of Use

Remove the direct probe and cable from the instrument and fit the high voltage probe and cable in its place. Set the facility selector to "+ or "-" VOLTS as required, and the range selector to a position considerably higher than the estimated value of voltage source.  More accurate readings may be made with successive settings of the range selector after taking the first trial measurement.

IMPCRTANT;

The probe is provided with an earthed guard ring in front of the hand-piece, and it is important that the operator's hand be behind this guard ring at all times, so that in the event of a flash-over occurring, current will pass only to the latter.

13.

2. METHOD OF USE

2.1	Warning Notes

(i)	DO NOT USE THIS INTRUMENT FOR THE MEASUREMENT OF A.C. AT 		POWER OUTLETS, ETC., WHERE THE ACTIVE LEAD CANNOT BE 		DISTINGUISHED.

(The frame of the instrument is earthed via the 3-core mains cable, and                                  		application of 	the test prods in incorrect polarity would therefore short-circuit 		the supply.)

WHEN MASURING ANY OTHER A.C. OR D.C. VOLTAGE IN WHICH 		ONE SIDE 	IS EARTHED, BE CAREFUL TO APPLY THE EARTHY 		PROD TO THE EARTHED 	SIDE OF THE VOLTAGE.

(ii)	Before taking any voltage measurements, refer to the maximum input voltage                		ratings given in para. 1.4.3.   On no account should these values be exceeded.

(iii)  	Do not connect the EARTH cable to a point of high potential. (The connector to                          		which this cable is attached is connected to the instrument case when measuring 		volts.)  Reversal of D.C. polarity is obtained by switching between ~'+" and "-"

VOLTS positions on the facility switch.

(iv)	When measuring the resistance of certain devices such as thermocouples, meters,                  		etc., a higher range than R x 1 or R x l0 should be used, because on these 1ower 		ranges the instrument applies up to l.5V. (depending on the resistance of the 		device being measured) in series with the external circuit, and could cause               		damage to the thermocouple or meter.

(v)	Exercise care when using the instrument as a milliammeter. The electronic              		protection devices incorporated for voltage applications are switched out of                   		circuit when the facility selector is set to "mA." position.

(vi)	Exercise extreme care when working on equipment in which high voltages 		appear.  Where possible it is advisable to follow the precautions listed below.

(a)	Remove the power from the equipment under test before connecting the                      			instrument.

(b)   Discharge large capacitors before making adjustments.

(c)   Check for breaks in the insulation wiring.

(d)	Disconnect test leads immediately a test is complete.

14.

(e)	Where practicable, use only one hand to make an adjustment or                   		measurement;  keep the other away from any part of the equipment or 		earth point.

(f)	Work in the presence of a second person who can render assistance in the               			case of accident.

2.2	Setting Up  Procedure

(i)	THE l.5V. CELL USED ON THE OHMS FACILITY IS NOT NORMALLY                     		FITTED IN THE INSTRUMENT WHEN SUPPLIED.  TO FIT THIS CELL, 		REMOVE THE INSTRUMENT FROM ITS CASE BY UNDOING THE                    		FRONT COEVER SCREWS, FIT THE CELL INTO THE APPROPRIATE 		CLAMP (UPPER RIGHT, WHEN VIEWED FROM THE FRONT), AND 		SOLDER/CONNECT (DEPENDING ON THE TYPE CF BATTERY USED) 		THE RED AND BLACK LEADS TO THE POSITIVE (CENTRE) AND                     		NEGATIVE (CELL CASE) POLES RESPECTIVELY.

(ii)	Identify the EARTH and OHMS cable and the Direct Probe (refer to Fig. 8) and                  		connect respectively to EARTH, OHMS and VOLTS (A.C.-D.C.) connectors on 		the front panel.

(iii)	Check the mechanical zero of the meter, and if necessary, re-adjust.   Plug into 		the mains and switch on.  Allow a short warm-up period.  Set facility selector to  		"+" VOLTS position.

(v)	Turn the ZERO ADJ. control until the meter needle coincides with the left-hand 		"0" on the scale.  Check the adjustment by now setting the facility switch to "-" 		VOLTS position.   If the meter needle moves from "0" again make an adjustment 		to. the ZERO ADJ. control.   If necessary, make alternate adjustments to the 		facility switch and ZERO ADJ. control until a setting is obtained which allows 		the facility switch to be set to either "+" or "-"  VOLTS without affecting the                      		needle.

(v)	Set the facility selector to the OHMS POSITION. The needle should be                 		deflected to  approximately full scale. Turn the OHMS ADJ. control until the          		needle 	coincides with the right hand edge of the OHMS scale.

The instrument is now ready for use.

2.3 	D.C. Voltage Measurement (using D.C. and Direct Probes)

Use the D.C. probe (type IR56020, indicated on Fig. 8) in conjunction with the direct 	probe R56020 for all D.C. voltage measurements.  Fit the probe piece over the tip of 	the direct probe and clip into place.

(i)	Set the facility selector to "+" VOLTS or "-" VOLTS as required (refer to               		warning note (iii) in 2.l).

15.

(ii)	Connect the EARTH cable to earth or low side of the voltage source under         		measurement.

(iii)	Set the range selector to a position considerably higher than that at first required.  		Use the probe as a conventional multimeter test prod.   Switch to progressively                		lower settings of the range selector to obtain more readable deflections.

Read direct from the appropriate scales.

2.4	Centre-zero Readings

The centre-zero facility allows observations of positive or negative excursions in D.C.        	voltages without the necessity of re-setting the facility selector to "-" or "+"  VOLTS, 	as the case may be.

(i)	Set the facility selector to "+" VOLTS.

(ii)	Turn the ZERO ADJ. control until the meter needle coincides 				with the centre "0" of the scale.

(iii)	Set the range selector to a position at least double the voltage to 				be measured.

(iv)	Take several test readings, re-setting the range selector for more 			convenient readings.

2.5	AC. Voltage Measurement (With Direct Probe alone)

Use the direct probe (type R56020) alone.  Low values  of R.M.S. (below 1.5V.) and 	peak-to-peak (below 4. 2V.) should be read on the separate LO-A.C. scales.

Higher voltages should be read on the "D.C. or R. M. S." scale graduated to 1.5 or 5 	respectively, in association with the appropriate multiplying factor given by the 	range               	selector.  Peak-to-peak values can be read simultaneously with the R.M.S. 	values.

(i)	Set the facility selector to A.C. VOLTS position and if necessary, adjust 			the ZERO ADJ. control to make the needle coincide with the left-hand           			"0"  of the scale.

(ii)	Set the range selector to a position considerably higher than that at first 			required.  Use the probe as a conventional test prod.   Switch to                               			progressively lower settings of the range selector to obtain more               			convenient readings.  Read direct from the appropriate scales.

Note that a delay will occur between the removal of the test prods from the source                 	voltage and the return of the meter needle to zero when measuring A.C. voltage.

16.

This delay is caused by the effects of circuit constants within the instrument, such constants having been chosen to enable more accurate measurements of pulse waveforms having low repetition rate.

2.6	 Resistance Measurement

2.6.1	Normal  Measurement of Resistance up to 1000 M.ohm

(i)	Refer to Warning Note (iv) of 2.1.

(ii)	Ensure that there is no voltage present in the circuit in which               			the instrument is to be used.

(iii)	Set the facility selector to OHMS position. Set the range         				selector to a convenient position (R X l000 being suitable)                      			connect the  OHMS PROBE to earth cable and turn ZERO ADJ                 			control until the meter needle coincides with the left hand "0"           			on the scale.

(v)	Disconnect probe from earth cable and check that meter needle          			is reading full scale.  If necessary readjust by means of OHMS ADJ 			control. The instrument is now ready for use and does not require new                              			setting up if  the range selector is set to a different 	position.

IF ONE END OF THE RESISTANCE TO BE MEASURED IS EARTHED,                        		OBSERVE THIS POLARITY WHEN USING 	THE OHMS PROBE                     		AND EARTH CABLE 	RESPECTIVELY.	CONNECT THE                            		LATTER TO THE 	EARTHED END OF THE 	RESISTOR, AND                   		APPLY THE 	PROBE TO THE OTHER.

(vi)	Multiply the reading on the ohms scale by the factor against the               			position of the range selector.

2.6.2		High Resistance (insulation) Measurement

The instrument has application as a high resistance (above 1,000 M.ohms) test                		unit when used in conjunction with an   external D.C. voltage (between 20V.                    		and 500V.), in a circuit as indicated in Fig. 4 below.

17.

(i)	Set the facility selector to "+" VOLTS, and measure the volts                               			at point "A".  Let this applied voltage be "V".

(ii)	Measure the voltage at point "B". Let this voltage be "v".

(iii) 	Calculate the unknown resistance from:-

R(in M.ohms)	11 (V-v)

v

2.7	Current Measurement

(i)   	Set the facility switch to "mA." position.

Observe the warning note (v) of 2e1.

(ii)	Use the instrument as a conventional milliammeter.

Connect the OHMS cable in the "+ side.  Set the range selector to a considerably higher position than that at first required, then make progressive adjustments for more convenient readings.  Read direct from the appropriate scale.

2.8	Measurement and Calculation of dbm

The chart given in Fig. 5 can be used in conjunction with the instrument to determine                                                              		dbm values corresponding to R.M..S. A.C. voltage values across a 	600-ohm load.                  		The dbm is defined as the number of decibels above or below a reference level of 1                  		milliwatt in 600 ohms at 1000 cycles.  Accordingly, 0 dbm  indicates a power of                      		1mW.; 10 dbm, 10 mW.;  20 dbm, 100 mW. etc.

Since dbm are defined with respect to a 600-ohm circuit, the power levels in                           	these circuits correspond to voltage values, thus allowing dbm to be measured in 	terms of R.M.S. voltages.   The chart given in Fig. 5 is drawn for application to 	600 ohm circuits, but may be used to convert R.M.S. to dbm values in load               	resistances other than 600 ohm, by the algebraic addition of a suitable increment 	(ΔDBM).

Several values of this increment have been worked out resistances commonly 		encountered, and are given in the table below. For other values of ΔDBM, use 		the formula given in equation (i).

18.

Resistive load at					ΔDBM (to be added to

1000 cycles					            the dbm value given on

the chart for 600 ohms)

600 ohms								0

500 ohms								+0.8

300 ohms								+3.0

250 ohms								+3.8

150 ohms								+6.0

50 ohms								+10.8

15 ohms								+16.0

8 ohms									+18.8

3.2 ohms								+22.7

Note that for load values greater than 600 ohms, ΔDBM will be a negative quantity, as indicated in the formula:-

ΔDBM = 10 log. 600   …….  (i)

R

Where R is the value of the resistance concerned.

19.

3. MAINTENANCE

3.1	Valve Replacement

If at any time it becomes necessary to replace either of the type 6AL5 	(alternative D77) or 12AU7 valves, they should be "aged" before installation, as              	follows:-

(a) 	Type 6AL5

Run the valve with heater voltage only applied for a period of 48 hours before 		placing it in the instrument.

(b)	Type 12AU7

Connect together the grids and cathode, apply the normal heater voltage and, 		in addition, an A.C. voltage of approximately 100V., 50-cycle between anode 		and cathode.  Allow to run for 24 hours.

3.2	Checking the Condition of the 1.5V. cell

(i)	Set the facility to OHMS.

(ii)	Set the range selector to R x 1.

(iii)	Rotate the OHMS ADJ. control for full-scale deflection on the meter.

(IV)	Short together the ohms and earth cables for about ten seconds.

(v)	Open the above short-circuit and observe the meter needle.  If there is 	appreciable deviation from F.S.D., a weak cell is indicated and should be 	replaced.

Do not allow exhausted cells to remain in the instrument and deteriorate 	because chemical action resulting therefore can cause permanent damage.

3.3	Electrical Balance Check

(i)	Set the facility selector to "+" VOLTS.

(ii)	Rotate the ZERO ADJ. control and check that it is possible to set the meter                                	needle at 60% of F.S.D. on any range.

(iii)	Set the facility selector to "-" VOLTS, and again turn the ZERO ADJ. control.                   	It should be possible to set the meter needle at 10% of F.S.D. on any range.

20.

(iv)	If the above results cannot be obtained, it is an indication of unbalance in the 		triode units of the type 12AU7 valve, and the latter should be replaced  (refer 		to 3.1)

3.4 	Calibration

If it becomes necessary to restore the initial accuracy of the instrument, i.e. to                     	recalibrate, the following procedure should be carried out:-

Fig. 6 indicates the relevant positions of the adjustable controls used in the 	calibration.

Note that the accuracy of calibration can only be comparable with the 	accuracy of  the voltage standard used in the procedure.

3.4.1	D.C. Calibration

(i)	 Check the mechanical zero of the meter, and if necessary, re-adjust

(ii)	Set the facility selector to "+" VOLTS.  Switch on the instrument and allow a              		warm-up period.  Check the mains voltage. The instrument should be                 		calibrated when operating from a steady 240V., 50-cycle supply.

21.

(iii)	Using the ZERO ADJ. control, set the meter needle to "0", and leave the                             		control set at this position throughout the remainder of the procedure.

(iv)	Set the range selector to the 0-50V. range.

(v)	Connect the earth cable to the negative terminal of an accurately adjusted                          		50V. D.C. source, and the D.C. probe (fitted over the direct probe) to the                   		positive terminal of this source.

(vi)	Adjust the "+"  D.C. Cal. Control (R27) to bring the pointer exactly to the                   		50V. mark on the "D.C.  or R.M.S." scale.

(vii)	Leaving the other conditions set, reverse the connections to the 50V. source             		and set the facility selector to the "-" VOLTS position.

(viii)	Check the other D.C. ranges against reliable voltage sources, making                        		allowances for the tolerance given in the specification.  If necessary re-                    		calibrate on a range other than 0-50V., and check all components associated 		with the facility.

3.4.2	A.C. Calibration

(i)	Carry out steps (i) to (iii) of 3.4.1

(ii)	Remove the D.C. probe from the direct probe.  Set the facility selector to A.C.                      		VOLTS, and range selector to 0-1.5V.

(iii)	Short together the direct probe and earth cable, and adjust the A.C. ZERO                          		control (R37) for a zero on the meter.  If it is not possible to zero the needle,                		interchange the two type 6AL5 valves, and if a zero is still unobtainable,           		replace one or both of the latter valves.

(iv)	Set the range selector to 0-50V.

(v)	Connect the earth cable and tip of the direct probe to an accurate 50V. , 50                  		cycle source, and check that the reading is within the tolerance stated in the               		electrical specification.  If outside this tolerance, check the components                                      		associated with the facility.

3.4.3	A.C. Compensation Adjustment (C2)

(i)	Set the facility selector to A.C. VOLTS.  Check that the meter is positioned at                    		zero, firstly by setting the mechanical zero, and later (after switching on) by                		adjusting the ZERO ADJ. control.

22.

(ii)	Set the range selector to 0-500V. and apply an accurate 150V., 100kc. Signal                            		to the instrument using the direct probe (alone) and earth cable.

(iii)	Adjust the variable air trimmer (C2) at the rear of the instrument for a 150V.                  		reading on the "D.C. or R.M.S." scale of the meter.

23.

4.   COMPONENT SCHEDULE.

When ordering replacement parts please quote ALL details given below for a particular component.

Circ.  Ref.                                     				A.W.A. Type Number

No.	     			Description			(unless otherwise stated)

(a)	Capacitors

C1		0.1 µf., 600v.w., paper

C2		2-18 µµf., 400V.W., variable air trimmer,	18836

C3		0.05 µf 400V.W., paper

C4		0.05 µf 400V.W., paper

C5		3,300 µµf., 500V.W., silvered ceramic	T.C.C. type CTH315		C6		3,300 µµf., 500V.W., silvered ceramic	T.C.C. type CTH315

C7		16 µf.,  350V.P., electrolytic			Ducon type ET1043

C8		25µµf., 500V.W., mica			Simplex  PT or SM

C9		3,300 µµfd., 500V.W., silvered ceramic	T.C.C. type CTH315

(b)	Connectors

CM1		Single-point, female, coaxial			C102-M

CF1		Single-point, female, coaxial			S55178

CF2		Single-point, female, coaxial			S55178

(c)	Meter

M1		Moving coil, movement 0-200 µA. F.S.D.,

Master S34 square, fitted with knife-edge

Pointer.  Scaled to Drg. 56010V7		464047

(d)	Resistors

R1	0.9 M.ohm ± 1%, 1W.,  carbon			Welwyn type A3634

R2	0.325 M.ohm ± 1%, ½ W., carbon insulated		Erie type 100

R3	0.15 M.ohm ± 1%, ¼ W., carbon, high-stability	Erie type 108

R4	20.3 M.0hm ± 2% 1W., carbon			Welwyn type A3635

R5	7 M.ohm ± 1W., carbon, high stability		Erie type A3634

R6	2 M.ohm ± ½ W., carbon, insulated			Erie type 100

R7	0.7 M.ohm ± 1%, ¼ W., Carbon, high-stability	Erie type 108

R8	0.2 M.ohm 1%, ¼ W., carbon, high-stability		Erie type 108

24.

Circ.  Ref.                                     				A.W.A. Type Number

No.	     			Description			(unless otherwise stated)

R9	70,000 ohm ± 1%, ¼ W., carbon, high stability	Erie type 108

R10	20,000 ohm ± 1%, ¼ W., carbon, high stability	Erie type 108

R11	10,000 ohm ± 1%, ¼ W., carbon, high stability	Erie type 108

R12	90 ohm ± 1%, ¼ W., carbon, high stability		Erie type 108

R13	9.5 ohm ± 1%, wire wound				56010V30

R14	900 ohm ± 1%, ¼ W., carbon, high stability	 	Erie type 108

R15	9,000 ohm ± 1%, ¼ W., carbon, high stability	Erie type 108

R16	90,000 ohm ± 1%, ¼ W., carbon, high stability	Erie type 108

R17	0.9 M. ohm ± 1%, ¼ W., carbon, high stability	Erie type 108

R18	9 M.ohm, ± 1%, 2W., carbon				Welwyn type A3635

R19	3.3 M. ohm ± 1%, ½  W., carbon, insulated		I.R.C. type BT½

R20	39,000 ohm ± 5% ½ W., carbon, insulated		I.R.C. type BT½

R21	20,000 ohm, variable potentiometer, carbon;

Less switch.  Dim. A, B and C respectively

¾” x 0.218” x ⅜”					I.R.C. type C.S.

R22	33,000 ohm ± 5% ½ W., carbon, insulated		I.R.C. type BT½

R23	1,800 ohm ± 5% ½ W., carbon, insulated		I.R.C. type BT½

R24	1,800 ohm ± 5% ½ W., carbon, insulated		I.R.C. type BT½

R25	47,000 ohm ± 5% ½ W., carbon, high stability	Erie type 109

R26	20,000 ohm, variable potentiometer, carbon;

Less switch.  Dim. A, B and C respectively

¾” x 0.218” x ⅜”					I.R.C. type C.S.

R27	10,000 ohm, variable potentiometer, carbon;

Less switch. Dim. A, 7/16”; screw driver

Slotted							I.R.C. type C.S.

R28	10,000 ohm, variable potentiometer, carbon;

Less switch. Dim. A, 7/16”; screw driver

Slotted							I.R.C. type C.S.

R29	1,800 ohm ± 5% ½ W., carbon, insulated		I.R.C. type BT½

R30	56,000 ohm ± 5% ½ W., carbon, insulated		I.R.C. type BT½

R31	47,000 ohm ± 5% ½ W., carbon, insulated		I.R.C. type BT½

R32	6.8 M. ohm ± 5% ½ W., carbon, insulated		I.R.C. type BT½

R33	1.8M. ohm ± 5% ½ W., carbon, insulated		I.R.C. type BT½

R34	0.68 M. ohm ± 5% ½ W., carbon, insulated		I.R.C. type BT½

R35	0.18 M. ohm ± 5% ½ W., carbon, insulated		I.R.C. type BT½

R36	0.1 M. ohm ± 5% ½ W., carbon, insulated		I.R.C. type BT½

R37	2 M. ohm, variable potentiometer; log or

Linear							I.R.C. type C.S.

or Morganite type A

R38	107.8 ohm, ± 1%, wire-wound			56010V109

R39	30.8 ohm, ± 1%, wire-wound				56010V109-1

R40	10.78 30.8 ohm, ± 1%, wire-wound			56010V109-2

R41	30.8 ohm, ± 1%, wire-wound				56010V109-3

R42	1.078 ohm, ± 1%, wire-wound			56010V109-4

R43	0.308 ohm, ± 1%, wire-wound			56010CT114

R44	0.154 ohm, ± 1%, wire-wound			56010CT115

25.

Circ.  Ref.                                     				A.W.A. Type Number

No.	     			Description			(unless otherwise stated)

(e)	Sockets

V1		7-Pin Miniature					19965

V2		7-Pin Miniature					19965

V3		9-Pin Miniature					code No. 793037

(f)	Switches

S1		Oak “H” type, rotary wafer				56010V13

S2		Oak “H” type, rotary wafer				56010V14

(g)	Transformer

T1		Power Transformer					2TU56576

(h)	Crystal Diode Probe

C1		500 µµf., ± 20% mica					T.C.C. type

CM30

R1		5.1 M.ohm ± 5%, 1W., carbon			I.R.C. B.T.A.

W1		Rectifier						G.E.X. 55-1