m

HEWLETT* PACKARD

Fundamentals of RF and Microwave Power Measurements

Application Note 64-1A

Application Note 64-1A

Fundamentals of RF and Microwave Power Measurements

Table of Contents

I. Introduction

The Importance of Power ..........................................................        1

A Brief History of Power Measurement      ..............................................        2

II. Power Measurement

Units and Definitions        ............................................................        4

Three Methods of Sensing Power      ...................................................        8

Key Power Sensor Parameters      .....................................................        8

The Hierarchy of Power Measurement, National Standards and Traceability      ...............        9

A New Sensor for Power Reference Transfer      ..........................................      11

III. Thermistor Sensors and Instrumentation

Thermistor Sensors      ..............................................................      12

Coaxial Thermistor Sensors      .......................................................      13

Waveguide Thermistor Sensors .....................................................      13

Bridges, from Wheatstone to Dual-Compensated DC Types ..............................      13

Thermistors as Power Transfer Standards ............................................      15

Other DC-Substitution Meters      .....................................................      15

IV. Thermocouple Sensors and Instrumentation

Principles of Thermocouples    .......................................................        17

The Thermocouple Sensor .........................................................        18

Power Meters for Thermocouple Sensors .............................................        22

Reference Oscillator ..............................................................        24

HP EPM Series Power Meters ......................................................        24

V. Diode Sensors and Instrumentation

Diode Detector Principles      .........................................................        26

Using Diodes for Sensing Power        ...................................................        28

New Wide-Dynamic-Range CW-Only Power Sensors      ...................................        30

A New Versatile Power Meter to Exploit 90 dB Range Sensors      ...........................        31

Traceable Power Reference    ........................................................        33

Signal Waveform Effects on the Measurement Uncertainty of Diode Sensors ................        34

VI. Measurement Uncertainty

Power Transfer, Generators and Loads ...............................................        37

RF Circuit Descriptions        ..........................................................        37

Reflection Coefficient .............................................................        39

Signal Flowgraph Visualization      ....................................................        40

Mismatch Uncertainty        ...........................................................        44

Mismatch Loss ..................................................................        45

Other Sensor Uncertainties      .......................................................        45

Calibration Factor    ...............................................................        46

Power Meter Instrumentation Uncertainties ..........................................        47

Calculating Total Uncertainty      .....................................................        49

Power Measurement Equation      .....................................................        50

Worst Case Uncertainty ...........................................................        52

RSS Uncertainty      ................................................................        52

New Method of Combining Power Meter Uncertainties      .................................        54

Power Measurement Model for ISO Process ...........................................        55

Standard Uncertainty of Mismatch Model ............................................        57

Example of Calculation of Uncertainty Using ISO Model ................................        58

VII. Power Measurement Instrumentation Compared

Accuracy vs. Power Level      .........................................................        62

Frequency Range and SWR (Reflection Coefficient)      ....................................        64

Speed of Response      ...............................................................        65

Automated Power Measurement      ...................................................        66

Susceptibility to Overload .........................................................        67

Signal Waveform Effects    ..........................................................        68

VIII. Peak Power Measurements

A Brief History of Peak Power Measurements      ........................................        69

Peak Power Analyzers ............................................................        69

IEEE Video Pulse Standards Adapted for Microwave Pulses .............................        71

Peak Power Waveform Definitions    ..................................................        72

Measuring Complex Waveforms other than Pulsed Power ...............................        73

Glossary and List of Symbols      ......................................................        75

Pulse terms and definitions, Figures 8-3 and 8-4, reprinted from IEEE STD 194-1977 and ANSI/IEEE STD 181-1977, Copyright

©

1977 by the Institute of Electrical and Electronics Engineers, Inc. The IEEE disclaims any responsibility or liability resulting

from the placement and use in this publication. Information is reprinted with the permission of the IEEE.

1

I. Introduction

This application note, AN64-1A, is a major revision of the 1977 edition of AN64-1, which has served for many years as a key reference for RF and microwave power measurement. It was written for two purposes:

1) to retain some of the original text of the fundamentals of RF and microwave power measurements, which tends to be timeless, and

2) to present more modern power measurement techniques and test equipment which represents the current state-of-the-art.

This note reviews the popular techniques and instruments used for measuring power, discusses error mechanisms, and gives principles for calculating overall measurement uncertainty. It describes metrology-oriented issues, such as the basic national standards, round robin intercomparisons and traceability processes. These will help users to establish an unbroken chain of calibration actions from the NIST (U.S. National Institute of Standards and Technology) or other national standard bodies, down to the final measurement setup on a production test line or a communication tower at a remote mountaintop. This note also discusses new measurement uncertainty processes such as the new ISO Guide to the Expression of Uncertainties in Measurement, and the USA version, ANSI/NCSL 540Z-2-1996, U.S. Guide for Expression of Uncertainty in Measurement, which defines new approaches to handling uncertainty calculations.

This introductory chapter reviews the importance of power quantities. Chapter II discusses units, defines terms such as average power and pulse power, reviews key sensors and their parameters, and overviews the hierarchy of power standards and the path of traceability to the United States National Reference Standard. Chapters III, IV, and V detail instrumentation for measuring power with the three most popular power sensing devices: thermistors, thermocouples, and diode detectors. Chapter VI covers power transfer, signal flowgraph analysis and mismatch uncertainty, along with the remaining uncertainties of power instrumentation and the calculation of overall uncertainty. Chapter VII compares the three popular methods for measuring average power. Peak and pulse power measurement and measurement of signals with complex modulations are discussed in Chapter VIII.

The Importance of Power

A system’s output power level is frequently the critical factor in the design, and ultimately the purchase and performance of almost all radio frequency and microwave equipment. The first key factor is the concept of equity in trade. When a customer purchases a product with specified power performance for a negotiated price, the final production-line test results need to agree with the customer’s incoming inspection data. These receiving, installation or commissioning phases, often occur at different locations, and sometimes across national borders. The various measurements must be consistent within acceptable uncertainties.

Secondly, measurement uncertainties cause ambiguities in realizable performance of a transmitter. For example, a ten-watt transmitter costs more than a five-watt transmitter. Twice the power output means twice the geographical area is covered or 40 percent more radial range for a communication system. Yet, if the overall measurement uncertainty of the final product test is on the order of ±0.5 dB, the unit actually shipped could have output power as much as 10% lower than the customer expects, with resulting lower headroom in its operating profiles.

2

Because signal power level is so important to the overall system performance, it is also critical when specifying the components that build up the system. Each component of a signal chain must receive the proper signal level from the previous component and pass the proper level to the succeeding component. Power is so important that it is frequently measured twice at each level, once by the vendor and again at the incoming inspection stations before beginning the next assembly level.

It is at the higher operating power levels where each decibel increase in power level becomes more costly in terms of complexity of design, expense of active devices, skill in manufacture, difficulty of testing, and degree of reliability. The increased cost per dB of power level is especially true at microwave frequencies, where the high-power solid state devices are inherently more costly and the guard-bands designed into the circuits to avoid maximum device stress are also quite costly.

Many systems are continuously monitored for output power during ordinary operation. This large number of power measurements and their importance dictates that the measurement equipment and techniques be accurate, repeatable, traceable, and convenient. The goal of this HP application note, and others, is to guide the reader in making those measurement qualities routine.

Because many of the examples cited above used the term “signal level,” the natural tendency might be to suggest measuring voltage instead of power. At low frequencies, below about 100 kHz, power is usually calculated from voltage measurements across a known impedance. As the frequency increases, the impedance has large variations, so power measurements become more popular, and voltage or current are calculated parameters.

At frequencies from about 30 MHz on up through the optical spectrum, the direct measurement of power is more accurate and easier. Another example of decreased usefulness is in waveguide transmission configurations where voltage and current conditions are more difficult to define.

A Brief History of Power Measurement

From the earliest design and application of RF and microwave systems, it was necessary to determine the level of power output. Some of the techniques were quite primitive by today’s standards. For example, when Sigurd and Russell Varian, the inventors of the klystron microwave power tube in the late 1930s, were in the early experimental stages of their klystron cavity, the detection diodes of the day were not adequate for those microwave frequencies. The story is told that Russell cleverly drilled a small hole at the appropriate position in the klystron cavity wall, and positioned a fluorescent screen alongside. This technique was adequate to reveal whether the cavity was in oscillation and to give a gross indication of power level changes as various drive conditions were adjusted.

Early measurements of high power system signals were accomplished by arranging to absorb the bulk of the system power into some sort of termination and measuring the heat buildup versus time. A simple example used for high power radar systems was the water-flow calorimeter. These were made by fabricating a glass or low-dielectric-loss tube through the sidewall of the waveguide at a shallow angle. Since the water was an excellent absorber of the microwave energy, the power measurement required only a measurement of the heat rise of the water from input to output, and a measure of the volumetric flow versus time. The useful part

3

of that technique was that the water flow also carried off the considerable heat from the source under test at the same time it was measuring the desired parameter.

Going into World War II, as detection crystal technology grew from the early galena cat-whiskers, detectors became more rugged and performed to higher RF and microwave frequencies. They were better matched to transmission lines, and by using comparison techniques with sensitive detectors, unknown microwave power could be measured against known values of power generated by calibrated signal generators.

Power substitution methods emerged with the advent of sensing elements which were designed to couple transmission line power into the sensing element.1 Barretters were positive-temperature-coefficient elements, typically metallic fuses, but they were frustratingly fragile and easy to burn out. Thermistor sensors exhibited a negative temperature coefficient and were much more rugged. By including such sensing elements as one arm of a 4-arm balanced bridge, DC or low-frequency AC power could be withdrawn as RF/MW power was applied, maintaining the bridge balance and yielding a substitution value of power.2

Commercial calorimeters had a place in early measurements. Dry calorimeters absorbed system power and by measurement of heat rise versus time, were able to determine system power. The 1960’s HP 434A power meter was an oil-flow calorimeter, with a 10 watt top range, and also used a heat comparison between the RF load and another identical load driven by DC power.3 Water-flow calorimeters were offered for medium to high power levels.

This application note will allot most of its space to the more modern, convenient and wider dynamic range sensor technologies which have developed since those early days of RF and microwave. Yet, it is hoped that some appreciation will be reserved for those early developers in this field for having endured the inconvenience and primitive equipment of those times. _________

1.    B.P. Hand, “Direct Reading UHF Power Measurement,” Hewlett-Packard Journal, Vol. 1, No. 59 (May, 1950).

2.      E.L. Ginzton, “Microwave Measurements,” McGraw-Hill, Inc., 1957.

3.      B.P. Hand, “An Automatic DC to X-Band Power Meter for the Medium Power Range,” Hewlett-Packard Journal, Vol. 9, No. 12 (Aug., 1958).

4

II. Power Measurement

Units and Definitions

Watt

The International System of Units (SI) has established the watt (W) as the unit of power; one watt is one joule per second. Interestingly, electrical quantities do not even enter into this definition of power. In fact, other electrical units are derived from the watt. A volt is one watt per ampere. By the use of appropriate standard prefixes the watt becomes the kilowatt (1 kW = 103W), milliwatt (1 mW = 10-3W), microwatt (1 µW = 10-6W), nanowatt (1 nW = 10-9W), etc.

dB

In many cases, such as when measuring gain or attenuation, the ratio of two powers, or relative power, is frequently the desired quantity rather than absolute power. Relative power is the ratio of one power level, P, to some other level or reference level, Pref. The ratio is dimensionless because the units of both the numerator and denominator are watts. Relative power is usually expressed in decibels (dB)

The dB is defined by

(PP )

dB = 10 log10                                                                                                                          (2-1)

ref

The use of dB has two advantages. First, the range of numbers commonly used is more compact; for example +63 dB to -153 dB is more concise than 2 x 106 to 0.5 x 10-15. The second advantage is apparent when it is necessary to find the gain of several cascaded devices. Multiplication of numeric gain is then replaced by the addition of the power gain in dB for each device.

dBm

Popular usage has added another convenient unit, dBm. The formula for dBm is similar to (2-1) except the denominator, Pref is always one milliwatt:

(1 mP W)

dBm = 10 log10                                                                                                                      (2-2)

In this expression, P is expressed in milliwatts and is the only variable, so dBm is used as a measure of absolute power. An oscillator, for example, may be said to have a power output of 13 dBm. By solving for P in (2-2), the power output can also be expressed as 20 mW. So dBm means “dB above one milliwatt,” (no sign is assumed +) but a negative dBm is to be interpreted as “dB below one milliwatt.” The advantages of the term dBm parallel those for dB; it uses compact numbers and aids the use of addition instead of multiplication when cascading gains or losses in a transmission system.

Power

The term “average power” is very popular and is used in specifying almost all RF and microwave systems. The terms “pulse power” and “peak envelope power” are more pertinent to radar and navigation systems.

In elementary theory, power is said to be the product of voltage and current. But for an AC voltage cycle, this product V x I varies during the cycle as shown by curve p in Figure 2-1, according to a 2f relationship. From that example, a sinusoidal generator produces a sinusoidal current as expected, but the product of voltage and current has a DC term as well as a component at twice the generator frequency. The word “power,” as most commonly used, refers to that DC component of the power product.

&

---*"

e

J-

e      *R

Figure 2-1. The product of voltage and current, P, varies during the sinusoidal cycle.

DC Component

All the methods of measuring power to be discussed (except for one chapter on peak power measurement) use power sensors which, by averaging, respond to the DC component. Peak power instruments and sensors have time constants in the sub-microsecond region, allowing measurement of pulsed power waveforms.

The fundamental definition of power is energy per unit time. This corresponds with the definition of a watt as energy transfer at the rate of one joule per second. The important question to resolve is over what time is the energy transfer rate to be averaged when measuring or computing power? From Figure 2-1 it is clear that if a narrow time interval is shifted around within one cycle, varying answers for energy transfer rate are found. But at radio and microwave frequencies, such microscopic views of the voltage-current product are not common. For this application note, power is defined as the energy transfer per unit time averaged over many periods of the lowest frequency (RF or microwave) involved.

A more mathematical approach to power for a continuous wave (CW) is to

find the average height under the curve of P in Figure 2-1. Averaging is

done by finding the area under the curve, that is by integrating, and then

dividing by the length of time over which that area is taken. The length of

time should be an exact number of AC periods. The power of a CW signal

at frequency (l/T0) is:

nT0

1                  e    sin              2T π0 t)• ip sin (2T π0

P =

nT0

n 1T0e(2T π0 )                                  (2T π0            )

ep sin                          t      • ip sin                        + f          dt      (2-3)

where T0 is the AC period, ep and ip represent peak values of e and i, f is the phase angle between e and i, and n is the number of AC periods. This yields (for n = 1, 2, 3 . . .):

epip P =                          cos f

2

(2-4)

If the integration time is many AC periods long, then, whether n is a precise integer or not makes a vanishingly small difference. This result for large n is the basis of power measurement.

For sinusoidal signals, circuit theory shows the relationship between peak and rms values as:

ep = √ 2 Erms and ip = √ 2 Irms                                                      (2-5)

Using these in (2-4) yields the familiar expression for power: P = Erms • Irms cos f

(2-6)

5

0

6

Average Power

Average power, like the other power terms to be defined, places further restrictions on the averaging time than just “many periods of the highest frequency.” Average power means that the energy transfer rate is to be averaged over many periods of the lowest frequency involved. For a CW signal, the lowest frequency and highest frequency are the same, so average power and power are the same. For an amplitude modulated wave, the power must be averaged over many periods of the modulation component of the signal as well.

In a more mathematical sense, average power can be written as:

nT,

n 1T, e

Pavg =                                        e(t) • i(t)dt                                                                          (2-7)

0 where T, is the period of the lowest frequency component of e(t) and i(t).

The averaging time for average power sensors and meters is typically from several hundredths of a second to several seconds and therefore this process obtains the average of most common forms of amplitude modulation.

Pulse Power

For pulse power, the energy transfer rate is averaged over the pulse width, t . Pulse width t is considered to be the time between the 50 percent rise-time/falltime amplitude points.

Mathematically, pulse power is given by

t

p        1t    e

P=                            e(t) • i(t)dt                                                                                          (2-8)

0

By its very definition, pulse power averages out any aberrations in the pulse envelope such as overshoot or ringing. For this reason it is called pulse power and not peak power or peak pulse power as is done in many radar references. The terms peak power and peak pulse power are not used here for that reason. See Chapter VIII for more explanation of measurements of pulsed power.

The definition of pulse power has been extended since the early days of microwave to be:

P avg                                                                                                                          (2-9)

P= p

Duty Cycle

where duty cycle is the pulse width times the repetition frequency. This extended definition, which can be derived from (2-7) and (2-8) for rectangular pulses, allows calculation of pulse power from an average power measurement and the duty cycle.

For microwave systems which are designed for a fixed duty cycle, peak power is often calculated by use of the duty cycle calculation along with an average power sensor. One reason is that the instrumentation is less expensive, and in a technical sense, the averaging technique integrates all the pulse imperfections into the average.

7

Figure 2-2. Pulse power P p is averaged over the pulse width.

The evolution of highly sophisticated radar, electronic warfare and navigation systems, often based on complex pulsed and spread spectrum technology, has led to more sophisticated instrumentation for characterizing pulsed RF power. To present a more inclusive picture of all pulsed power measurements, Chapter VIII Peak Power Instrumentation is presented later in this note. Theory and practice and detailed waveform definitions are presented for those applications.

P„ —i

avg

| A | -*------ | T- 1 |  | Duty Cycle = xfr |  |
| --- | --- | --- | --- | --- | --- |
|  |  | ' f r |  |  |  |
|  |  |  |  |  | — . |

Peak Envelope Power

For certain more sophisticated microwave applications, and because of the need for greater accuracy, the concept of pulse power is not totally satisfactory. Difficulties arise when the pulse is intentionally non rectangular or when aberrations do not allow an accurate determination of pulse width x. Figure 2-3 shows an example of a Gaussian pulse shape, used in certain navigation systems, where pulse power, by either (2-8) or (2-9), does not give a true picture of power in the pulse.    Peak envelope power is a term for describing the maximum power. Envelope power will first be discussed.

Figure 2-3. A Gaussian pulse and the different kinds of power.

P p using (2-9) P p using (2-8)

Envelope power is measured by making the averaging time much less than 1/fm where fm is the maximum frequency component of the modulation waveform. The averaging time is therefore limited on both ends: (1) it must be small compared to the period of the highest modulation frequency, and (2) it must be large enough to be many RF cycles long.

By continuously displaying the envelope power on an oscilloscope, (using a detector operating in its square-law range), the oscilloscope trace will show the power profile of the pulse shape.    (Square law means the detected output voltage is proportional to the input RF power, that is the square of the input voltage.)    Peak envelope power, then, is the maximum value of the envelope power (see Figure 2-3).    For perfectly rectangular pulses, peak envelope power is equal to pulse power as defined above.    Peak power analyzers are specifically designed to completely characterize such waveforms. See Chapter VIII.

8

Average power, pulse power, and peak envelope power all yield the same answer for a CW signal. Of all power measurements, average power is the most frequently measured because of convenient measurement equipment with highly accurate and traceable specifications. Pulse power and peak envelope power can often be calculated from an average power measurement by knowing the duty cycle. Average power measurements therefore occupy the greatest portion of this application note series.

Other Waveforms and Multiple Signals

The recent explosion of new RF and microwave systems which depend on signal formats other than simple pulse or AM/FM modulation has made power measuring techniques more critical. Modern systems use fast digital phase-shift-keyed modulations, wide-channel, multiple carrier signals, and other complex formats which complicate selection of sensor types. In particular, the popular diode sensors are in demand because of their wide dynamic range. But the sophisticated shaping circuits need careful analysis when used in non-CW signal environments. More explanation is detailed in Chapter V.

Three Methods of Sensing Power

There are three popular devices for sensing and measuring average power at RF and microwave frequencies. Each of the methods uses a different kind of device to convert the RF power to a measurable DC or low frequency signal. The devices are the thermistor, the thermocouple, and the diode detector. Each of the next three chapters discusses in detail one of those devices and its associated instrumentation. Each method has some advantages and disadvantages over the others. After the individual measurement sensors are studied, the overall measurement errors are discussed in Chapter VI. Then the results of the three methods are summarized and compared in Chapter VII.

The general measurement technique for average power is to attach a properly calibrated sensor to the transmission line port at which the unknown power is to be measured. The output from the sensor is connected to an appropriate power meter. The RF power to the sensor is turned off and the power meter zeroed. This operation is often referred to as “zero setting” or “zeroing.” Power is then turned on. The sensor, reacting to the new input level, sends a signal to the power meter and the new meter reading is observed.

Key Power Sensor Parameters

In the ideal measurement case above, the power sensor absorbs all the power incident upon the sensor. There are two categories of non-ideal behavior that are discussed in detail in Chapters VI and VII, but will be introduced here.

First, there is likely an impedance mismatch between the characteristic

impedance of the RF source or transmission line and the RF input impedance

of the sensor. Thus, some of the power that is incident on the sensor is

reflected back toward the generator rather than dissipated in the sensor.

The relationship between incident power P , reflected power P , and dissipat-ir

ed power P , is: d

Pi = Pr + Pd                                                                                                                                  (2-10)

The relationship between Pi and Pr for a particular sensor is given by the

sensor reflection coefficient magnitude r      .

,

P= rP                                                                                                                                        (2-11)

r,2        i

9

Reflection coefficient magnitude is a very important specification for a

power sensor because it contributes to the most prevalent source of error,

mismatch uncertainty, which is discussed in Chapter VI. An ideal power

sensor has a reflection coefficient of zero, no mismatch. While a r          of 0.05

,

or 5 percent (equivalent to an SWR of approximately 1.11) is preferred for most situations, a 50 percent reflection coefficient would not be suitable for most situations due to the large measurement uncertainty it causes. Some early waveguide sensors were specified at a reflection coefficient of 0.35.

The second cause of non-ideal behavior is that RF power is dissipated in

places other than in the power sensing element. Only the actual power

dissipated in the sensor element gets metered. This effect is defined as

the sensor’s effective efficiency h      . An effective efficiency of 1 (100%)

e

means that all the power entering the sensor unit is absorbed by the sensing element and metered — no power is dissipated in conductors, sidewalls, or other components of the sensor.

Microcalorimeter

National Reference

Standard

NIST

Working Standards

Measurement Reference Standard

Transfer Standard

NIST

Commercial

Standards

Laboratory

Manufacturing Facility

General Test Equipment

User

Figure 2-4. The traceability path of power references from the United States National Reference Standard.

The most frequently used specification of a power sensor is called the calibration factor, Kb. Kb is a combination of reflection coefficient and effective efficiency according to

Kb = Te (1 - P      2)

(2-12)

If a sensor has a Kb of 0.90 (90%) the power meter would normally indicate a power level that is 10 percent lower than the incident power Pi. Most power meters have the ability to correct the lower-indicated reading by setting a calibration factor dial (or keyboard or HP-IB on digital meters) on the power meter to correspond with the calibration factor of the sensor at the frequency of measurement. Calibration factor correction is not capable of correcting for the total effect of reflection coefficient. There is still a mismatch uncertainty that is discussed in Chapter VI.

The Hierarchy of Power Measurements, National Standards and Traceability

Since power measurement has important commercial ramifications, it is important that power measurements can be duplicated at different times and at different places. This requires well-behaved equipment, good measurement technique, and common agreement as to what is the standard watt. The agreement in the United States is established by the National Institute of Standards and Technology (NIST) at Boulder, Colorado, which maintains a National Reference Standard in the form of various microwave microcalorimeters for different frequency bands. 1, 2    When a power sensor can be referenced back to that National Reference Standard, the measurement is said to be traceable to NIST.

The usual path of traceability for an ordinary power sensor is shown in Figure 2-4. At each echelon, at least one power standard is maintained for the frequency band of interest. That power sensor is periodically sent to the next higher echelon for recalibration, then returned to its original level. Recalibration intervals are established by observing the stability of a device between successive recalibrations. The process might start with recalibration every few months. Then, when the calibration is seen not to change, the interval can be extended to a year or so.

10

Figure 2-5. Schematic cross-section of the NIST coaxial microcalori-meter at Boulder, CO. The entire sensor configuration is maintained under a water bath with a highly-stable temperature so that RF to DC substitutions may be made precisely.

Each echelon along the traceability path adds some measurement uncertainty. Rigorous measurement assurance procedures are used at NIST because any error at that level must be included in the total uncertainty at every lower level. As a result, the cost of calibration tends to be greatest at NIST and reduces at each lower level. The measurement comparison technique for calibrating a power sensor against one at a higher echelon is discussed in other documents, especially those dealing with round robin procedures. 3,4

The National Power Reference Standard for the U.S. is a microcalorimeter maintained at the NIST in Boulder, CO, for the various coaxial and waveguide frequency bands offered in their measurement services program. These measurement services are described in NIST SP-250, available from NIST on request.5 They cover coaxial mounts from 10 MHz to 26.5 GHz and waveguide from 8.2 GHz to the high millimeter ranges of 96 GHz.

A microcalorimeter measures the effective efficiency of a DC substitution sensor which is then used as the transfer standard. Microcalorimeters operate on the principle that after applying an equivalence correction, both DC and absorbed microwave power generate the same heat. Comprehensive and exhaustive analysis is required to determine the equivalence correction and account for all possible thermal and RF errors, such as losses in the transmission lines and the effect of different thermal paths within the microcalorimeter and the transfer standard. The DC-substitution technique is used because the fundamental power measurement can then be based on DC voltage (or current) and resistance standards. The traceability path leads through the micro-calorimeter (for effective efficiency, a unit-less correction factor) and finally back to the national DC standards.

In addition to national measurement services, other industrial organizations often participate in comparison processes known as round robins (RR). A round robin provides measurement reference data to a participating lab at very low cost compared to primary calibration processes. For example, the National Conference of Standards Laboratories, a non-profit association of over 1400 world-wide organizations, maintains round robin projects for many measurement parameters, from dimensional to optical. The NCSL Measurement Comparison Committee oversees those programs.4

For RF power, a calibrated thermistor mount starts out at a “pivot lab,” usually one with overall RR responsibility, then travels to many other reference labs to be measured, returning to the pivot lab for closure of measured data. Such mobile comparisons are also carried out between National Laboratories of various countries as a routine procedure to assure international measurements at the highest level.

Microwave power measurement services are available from many National Laboratories around the world, such as the NPL in the United Kingdom and PTB in Germany. Calibration service organizations are numerous too, with names like NAMAS in the United Kingdom.

11

A New Sensor for Power Reference Transfer

Although thermistor sensors have served for decades as the primary portable sensor for transferring RF power, they have several drawbacks. Their frequency range is limited, and thermistor impedance matches were never as good as most comparison processes would have preferred. Except for the most extreme measurement cases requiring highest accuracy, few modern procedures call for coaxial or waveguide tuners to match out reflections from mismatched sensors at each frequency.

A new cooperative research effort between HP and NIST is aimed at producing a novel resistive sensor designed for the express purpose of transferring microwave power references. The element is called a resistive power sensor, and its DC to 50 GHz frequency range provides for DC substitution techniques in a single sensor in 2.4 mm coax. The technology is based on microwave microcircuit fabrication of a precise 50 Ω resistive element on Gallium-Arsenide. The resistor presents a positive temperature coefficient when heated, and can be operated by an NIST Type 4 Power Meter.

Future processes for reference power transfer will likely be based on such a new technology because of its wide frequency coverage and excellent SWR.

____________

1.    M.P. Weidman and P.A. Hudson, “WR-10 Millimeterwave Microcalorimeter,” NIST Technical Note 1044, June, 1981.

2.    F.R. Clague, “A Calibration Service for Coaxial Reference Standards for Microwave Power,” NIST Technical Note 1374, May, 1995.

3.    National Conference of Standards Laboratories, Measurement Comparison Committee, Suite 305B, 1800 30th St. Boulder, CO 80301.

4.    M.P. Weidman, “Direct Comparison Transfer of Microwave Power Sensor Calibration,” NIST Technical Note 1379, January, 1996.

5.    Special Publication 250; NIST Calibration Services, 1991 Edition.

General References

R.W. Beatty, “Intrinsic Attenuation,” IEEE Trans. on Microwave Theory and Techniques, Vol. I I, No. 3 (May, 1963) 179-182.

R.W. Beatty, “Insertion Loss Concepts,” Proc. of the IEEE. Vol. 52, No. 6 (June, 1966) 663-671.

S.F. Adam, “Microwave Theory & Applications,” Prentice-Hall, 1969.

C.G. Montgomery, “Technique of Microwave Measurements,” Massachusetts Institute of Technology, Radiation Laboratory Series, Vol. 11. McGraw-Hill, Inc., 1948.

Mason and Zimmerman. “Electronic Circuits, Signals and Systems,” John Wiley and Sons, Inc., 1960.

12

III. Thermistor Sensors and Instrumentation

Bolometer sensors, especially thermistors, have held an important historical position in RF/microwave power measurements. However, in recent years thermocouple and diode technologies have captured the bulk of those applications because of their increased sensitivities, wider dynamic ranges and higher power capabilities. Yet, thermistors are still the sensor of choice for certain applications, such as transfer standards, because of their power substitution capability. So, although this chapter is shortened from AN64-1, the remaining material should be adequate to understand the basic theory and operation of thermistor sensors and their associated dual-balanced bridge power meter instruments.

Bolometers are power sensors that operate by changing resistance due to a change in temperature. The change in temperature results from converting RF or microwave energy into heat within the bolometric element. There are two principle types of bolometers, barretters and thermistors. A barretter is a thin wire that has a positive temperature coefficient of resistance. Thermistors are semiconductors with a negative temperature coefficient.

To have a measurable change in resistance for a small amount of dissipated RF power, a barretter is constructed of a very thin and short piece of wire, or alternately, a 10 mA instrument fuse. The maximum power that can be measured is limited by the burnout level of the barretter, typically just over 10 mW, and they are seldom used anymore.

The thermistor sensor used for RF power measurement is a small bead of metallic oxides, typically 0.4 mm diameter with 0.03 mm diameter wire leads. Thermistor characteristics of resistance vs. power are highly nonlinear, and vary considerably from one thermistor to the next. Thus the balanced-bridge technique always maintains the thermistor element at a constant resistance, R, by means of DC or low frequency AC bias. As RF power is dissipated in the thermistor, tending to lower R, the bias power is withdrawn by just the proper amount to balance the bridge and keep R the same value. The decrease in bias power should be identical to the increase in RF power. That decrease in bias power is then displayed on a meter to indicate RF power.

Thermistor Sensors

Thermistor elements are mounted in either coaxial or waveguide structures so they are compatible with common transmission line systems used at microwave and RF frequencies. The thermistor and its mounting must be designed to satisfy several important requirements so that the thermistor element will absorb as much of the power incident on the mount as possible. First, the sensor must present a good impedance match to the transmission line over the specified frequency range. The sensor must also have low resistive and dielectric losses within the mounting structure because only power that is dissipated in the thermistor element can be registered on the meter. In addition, mechanical design must provide isolation from thermal and physical shock and must keep leakage small so that microwave power does not escape from the mount in a shunt path around the thermistor. Shielding is also important to prevent extraneous RF power from entering the mount.

Modern thermistor sensors have a second set of compensating thermistors to correct for ambient temperature variations. These compensating thermistors are matched in their temperature-resistance characteristics to the detecting thermistors. The thermistor mount is designed to maintain electrical isolation between the detecting and compensating thermistors yet keeping the thermistors in very close thermal contact.

13

Coaxial Thermistor Sensors

The HP 478A and 8478A thermistor mounts (thermistor mount was the earlier name for sensor) contain four matched thermistors, and measure power from 10 MHz to 10 and 18 GHz. The two RF-detecting thermistors, bridge-balanced to 100 Ω each, are connected in series (200 Ω) as far as the DC bridge circuits are concerned. For the RF circuit, the two thermistors appear to be connected in parallel, presenting a 50 Ω impedance to the test signal. The principle advantage of this connection scheme is that both RF thermistor leads to the bridge are at RF ground. See Figure 3-1 (a).

Thermal Conducting Block

|  |  |  |  | Compensation |  |
| --- | --- | --- | --- | --- | --- |
|  |  | Rc |  | Bridge Bias |  |
|  |  | —WSr-f | p |  |  |
|  | sRc ; |  | ;Cb |  |  |
| Cc |  |  |  |  |  |
|  |  | •Rd ' | ;Rd |  | RF Bridge |
| RF Power _ |  |  |  |  | Bias |
|  |  | iCb |  |  |  |
|  |  |  |  |  | I ° |

(a)

(R C ) Compensating

Thermistor

(Underneath)

Heat

Conductive

Strap

Thermal Isolation Disc

(b)

Figure 3-1.

(a) HP 478A coaxial sensor simplified diagram.

(b) HP 486A waveguide sensor construction.

Compensating thermistors, which monitor changes in ambient temperature but not changes in RF power, are also connected in series. These thermistors are also biased to a total of 200 Ω by a second bridge in the power meter, called the compensating bridge. The compensating thermistors are completely enclosed in a cavity for electrical isolation from the RF signal. But they are mounted on the same thermal conducting block as the detecting thermistors. The thermal mass of the block is large enough to prevent sudden temperature gradients between the thermistors. This isolates the system from thermal inputs such as human hand effects.

There is a particular error, called dual element error, that is limited to coaxial thermistor mounts where the two thermistors are in parallel for the RF energy, but in series for DC. If the two thermistors are not quite identical in resistance, then more RF current will flow in the one of least resistance, but more DC power will be dissipated in the one of greater resistance. The lack of equivalence in the dissipated DC and RF power is a minor source of error that is proportional to power level. For HP thermistor sensors, this error is less than 0.1 percent at the high power end of their measurement range and is therefore considered as insignificant in the error analysis of Chapter VI.

Waveguide Thermistor Sensors

The HP 486A-series of waveguide thermistor mounts covers frequencies from 8 to 40 GHz. See Figure 3-1 (b). Waveguide sensors up to 18 GHz utilize a post-and-bar mounting arrangement for the detecting thermistor. The HP 486A-series sensors covering the K and R waveguide band (18 to 26.5 GHz and 26.5 to 40 GHz) utilize smaller thermistor elements which are biased to an operating resistance of 200 Ω, rather than the 100 Ω used in lower frequency waveguide units. Power meters provide for selecting the proper 100 or 200 Ω bridge circuitry to match the thermistor sensor being used.

Bridges, from Wheatstone to Dual-Compensated DC Types

Over the decades, power bridges for monitoring and regulating power sensing thermistors have gone through a major evolution. Early bridges such as the simple Wheatstone type were manually balanced. Automatically-balanced bridges, such as the HP 430C of 1952, provided great improvements in convenience but still had limited dynamic range due to thermal drift on their 30 µW (full scale) range. In 1966, with the introduction of the first temperature-compensated meter, the HP 431A, drift was reduced so much that meaningful measurements could be made down to 1 µW.1

The HP 432A power meter, uses DC and not audio frequency power to maintain balance in both bridges. This eliminates earlier problems pertaining to the 10 kHz bridge drive signal applied to the thermistors. The HP 432A has the further convenience of an automatic zero set, eliminating the need for the operator to precisely reset zero for each measurement.

14

The HP 432A features an instrumentation accuracy of 1 percent. It also provides the ability to externally measure the internal bridge voltages with higher accuracy DC voltmeters, thus permitting a higher accuracy level for power transfer techniques to be used. In earlier bridges, small, thermo-electric voltages were present within the bridge circuits which ideally should have cancelled in the overall measurement. In practice, however, cancellation was not complete. In certain kinds of measurements this could cause an error of 0.3 µW. In the HP 432A, the thermo-electric voltages are so small, compared to the metered voltages, as to be insignificant.

The principal parts of the HP 432A (Figure 3-2) are two self-balancing bridges, the meter-logic section, and the auto-zero circuit. The RF bridge, which contains the detecting thermistor, is kept in balance by automatically varying the DC voltage Vrf, which drives that bridge. The compensating bridge, which contains the compensating thermistor, is kept in balance by automatically varying the DC voltage Vc, which drives that bridge.

The power meter is initially zero-set (by pushing the zero-set button) with no applied RF power by making Vc equal to Vrfo (Vrfo means Vrf with zero RF power). After zero-setting, if ambient temperature variations change thermistor resistance, both bridge circuits respond by applying the same new voltage to maintain balance.

Figure 3-2. Simplified diagram of the HP 432A power meter.

If RF power is applied to the detecting thermistor, Vrf decreases so that

P rf

V f      2  ro

4R

V f 2  r

4R

(3-1)

where Prf is the RF power applied and R is the value of the thermistor resistance at balance. But from zero-setting, Vrfo= Vc so that

P rf

1 4R

(V c 2 - V rf 2)

(3-2)

15

which can be written

1

Prf =                    (Vc – Vrf) (Vc + Vrf)                                                                    (3-3)

4R

The meter logic circuitry is designed to meter the voltage product shown in equation (3-3). Ambient temperature changes cause Vc and Vrf to change so there is zero change to Vc2 - Vrf2 and therefore no change to the indicated Prf.

As seen in Figure 3-2, some clever analog circuitry is used to accomplish the multiplication of voltages proportional to (Vc - Vrf) and (Vc + Vrf) by use of a voltage-to-time converter. In these days, such simple arithmetic would be performed by the ubiquitous micro-processor, but the HP 432A predated that technology, and performs well without it.

The principal sources of instrumentation uncertainty of the HP 432A lie in the metering logic circuits. But Vrf and Vc are both available at the rear panel of the HP 432A. With precision digital voltmeters and proper procedure, those outputs allow the instrumentation uncertainty to be reduced to ±0.2 percent for many measurements. The procedure is described in the operating manual for the HP 432A.

Thermistors as Power Transfer Standards

For special use as transfer standards, the U.S. National Institute for Standards and Technology (NIST), Boulder, CO, accepts thermistor mounts, both coaxial and waveguide, to transfer power parameters such as calibration factor, effective efficiency and reflection coefficient in their measurement services program. To provide those services below 100 MHz, NIST instructions require sensors specially designed for that performance.

One example of a special power calibration transfer is the one required to precisely calibrate the internal 50 MHz, 1 mW power standard in the HP 437B and 438A power meters, which use a family of thermocouple sensors. That internal power reference is needed since thermocouple sensors do not use the power substitution technique. For the power reference, a specially-modified HP 478A thermistor sensor with a larger coupling capacitor is available for operation from 1 MHz to 1 GHz. It is designated the HP H55 478A and features an SWR of 1.35 over its range. For an even lower transfer uncertainty at 50 MHz, the HP H55 478A can be selected for 1.05 SWR at 50 MHz. This selected model is designated the HP H75 478A.

HP H76 478A thermistor sensor is the H75 sensor which has been specially calibrated in the HP Microwave Standards Lab with a 50 MHz power reference traceable to NIST.

Other coaxial and waveguide thermistor sensors are available for metrology use.

Other DC-Substitution Meters

Other self-balancing power meters can also be used to drive HP thermistor sensors for measurement of power. In particular, the NIST Type 4 power meter, designed by the NIST for high-accuracy measurement of microwave power is well suited for the purpose. The Type 4 meter uses automatic balancing, along with a four-terminal connection to the thermistor sensor and external high precision DC voltage instrumentation. This permits lower uncertainty than standard power meters are designed to accomplish.

16

Conclusions

There are some advantages to thermistor power measurements that have not been obvious from the above discussion or from data sheet specifications. Thermistor mounts are the only present-day sensors which allow power substitution measurement techniques, and thus retain importance for traceability and absolute reference to national standards and DC voltages.

The fundamental premise in using a thermistor for power measurements is that the RF power absorbed by the thermistor has the same heating effect on the thermistor as the DC power. The measurement is said to be “closed loop,” because the feedback loop corrects for minor device irregularities.

1. R.F. Pramann, “A Microwave Power Meter with a Hundredfold Reduction in Thermal Drift,” Hewlett-Packard Journal, Vol. 12, No. 10 (June, 1961).

General References

“IEEE Standard Application Guide for Bolometric Power Meters,” IEEE

Std. 470-1972.

“IEEE Standard for Electrothermic Power Meters,” IEEE Std. 544-1976

17

IV. Thermocouple Sensors and Instrumentation

Thermocouple sensors have been the detection technology of choice for sensing RF and microwave power since their introduction in 1974. The two main reasons for this evolution are: 1) they exhibit higher sensitivity than previous thermistor technology, and 2) they feature an inherent square-law detection characteristic (input RF power is proportional to DC voltage out).

Since thermocouple are heat-based sensors, they are true “averaging detectors.” This recommends them for all types of signal formats from CW to complex digital phase modulations. In addition, they are more rugged than thermistors, make useable power measurements down to 0.3 µW (- 30 dBm, full scale), and have lower measurement uncertainty because of better SWR.

Figure 4-1. Heat at one end of a metal rod gives rise to an electric field.

The evolution to thermocouple technology is the result of combining thin-film and semiconductor technologies to give a thoroughly understood, accurate, rugged, and reproducible power sensor. This chapter briefly describes the principles of thermocouples, the construction and design of modern thermocouple sensors, and the instrumentation used to measure their rather tiny sensor DC-output levels.

Principles of Thermocouples

Thermocouples are based on the fact that dissimilar metals generate a voltage due to temperature differences at a hot and a cold junction of the two metals. As a simple example of the physics involved, imagine a long metal rod that is heated at the left end as in Figure 4-1. Because of the increased thermal agitation at the left end, many additional electrons become free from their parent atoms. The increased density of free electrons at the left causes diffusion toward the right. There is also a force attempting to diffuse the positive ions to the right but the ions are locked into the metallic structure and cannot migrate. So far, this explanation has not depended on Coulomb forces. The migration of electrons toward the right is by diffusion, the same physical phenomenon that tends to equalize the partial pressure of a gas throughout a space.

Each electron that migrates to the right leaves behind a positive ion. That ion tends to attract the electron back to the left with a force given by Coulomb’s law. The rod reaches equilibrium when the rightward force of heat-induced diffusion is exactly balanced by the leftward force of Coulomb’s law. The leftward force can be represented by an electric field pointing toward the right. The electric field, summed up along the length of the rod, gives rise to a voltage source called the Thomson emf. This explanation is greatly simplified but indicates the principle.

18

The same principles apply at a junction of dissimilar metals where different free electron densities in the two different metals give rise to diffusion and an emf. The name of this phenomenon is the Peltier effect.

Figure 4-2. Total thermocouple output is the resultant of several thermo-electrical voltages generated along the two-metal circuit.

A thermocouple is usually a loop or circuit of two different materials as shown in Figure 4-2. One junction of the metals is exposed to heat, the other is not. If the loop remains closed, current will flow in the loop as long as the two junctions remain at different temperatures. If the loop is broken to insert a sensitive voltmeter, it will measure the net emf. The thermocouple loop uses both the Thomson emf and the Peltier emf to produce the net thermoelectric voltage. The total effect is also known as the Seebeck emf.

L

~!

Diffused Region

I Tantalum Nitride I Ta 2 N

I          Web

I 0.005 mm

Figure 4-3. Photo-micrograph of the structure of the HP 8481A thermocouple chip on a thin silicon web.

Sometimes many pairs of junctions or thermocouples are connected in series and fabricated in such a way that the first junction of each pair is exposed to heat and the second is not. In this way the net emf produced by one thermocouple adds to that of the next, and the next, etc., yielding a larger thermoelectric output. Such a series connection of thermocouples is called a thermopile.

Early thermocouples for sensing RF power were frequently constructed of the metals bismuth and antimony. To heat one junction in the presence of RF energy, the energy was dissipated in a resistor constructed of the metals making up the junction. The metallic resistor needed to be small in length and cross section to form a resistance high enough to be a suitable termination for a transmission line. Yet, the junction needed to produce a measurable change in temperature for the minimum power to be detected and measured. Thin-film techniques were normally used to build metallic thermocouples. These small metallic thermocouples tended to have parasitic reactances and low burnout levels. Further, larger thermopiles, which did have better sensitivity, tended to be plagued by reactive effects at microwave frequencies because the device dimensions became too large for good impedance match at higher microwave frequencies.

The Thermocouple Sensor

The modern thermocouple sensor was introduced in 19741, and is exemplified by the HP 8481A power sensor. It was designed to take advantage of both semiconductor and microwave thin-film technologies. The device, shown in Figure 4-3, consists of two thermocouples on a single integrated-circuit chip. The main mass of material is silicon.

The principal structural element is the frame made of p-type silicon, which supports a thin web of n-type silicon. The smoothly sloped sides of the frame result from an anisotropic etch acting on the silicon crystal planes. The thin web is produced by epitaxially growing it on the p-type substrate and then suitably controlling the etch, which also reveals the surface of the diffused regions in the web.

0.81 mm

Go d Beam

Lead

19

Figure 4-4. Cross section of one thermocouple. Power dissipated in the tanta lum-nitride resistor heats the hot junction.

At the end of the diffused region near the center of the web, a second metal penetration to the web is made by a tantalum nitride film. This contact is the hot junction of the thermocouple. The tantalum nitride, which is deposited on the silicon oxide surface, continues to the edge of the frame, where it contacts the opposite beam lead terminal. This tantalum nitride forms the other leg of the thermocouple.

The other edge of the resistor and the far edge of the silicon chip have gold beam-lead contacts. The beam leads not only make electrical contact to the external circuits, but also provide mounting surfaces for attaching the chip to a substrate, and serve as good thermal paths for conducting heat away from the chip. This tantalum-nitride resistor is not at all fragile in contrast to similar terminations constructed of highly conductive metals like bismuth/antimony

As the resistor converts the RF energy into heat, the center of the chip, which is very thin, gets hotter than the outside edge for two reasons. First, the shape of the resistor causes the current density and the heat generated to be largest at the chip center. Second, the outside edges of the chip are thick and well cooled by conduction through the beam leads. Thus, there is a thermal gradient across the chip which gives rise to the thermoelectric emf. The hot junction is the resistor-silicon connection at the center of the chip. The cold junction is formed by the outside edges of the silicon chip between the gold and diffused silicon region.

The thin web is very important, because the thermocouple output is proportional to the temperature difference between the hot and cold junctions. In this case the web is fabricated to be 0.005 mm thick.    Silicon is quite a good thermal conductor, so the web must be very thin if reasonable temperature differences are to be obtained from low power inputs.

Silicon Oxide SiO2

Tantaium Nitride 2N

Cold Junction

Hot Junction

Figure 4-4 is a cross section through one of the thermocouples. One gold beam lead terminal penetrates the insulating silicon oxide surface layer to contact the web over the edge of the frame. This portion of the web has been more heavily doped by diffusing impurities into it. The connection between the gold lead and the diffused region is the cold junction of the thermocouple, and the diffused silicon region is one leg of the thermocouple.

20

Figure 4-5. Schematic diagram of the HP 8481A thermocouple power

The HP 8481A power sensor contains two identical thermocouples on one chip, electrically connected as in Figure 4-5. The thermocouples are connected in series as far as the DC voltmeter is concerned. For the RF input frequencies, the two thermocouples are in parallel, being driven through

coupling capacitor C . Half the RF current flows through each thermocouple. c

Each thin-film resistor and the silicon in series with it has a total resistance of 100 Ω. The two thermocouples in parallel form a 50 Ω termination to the RF transmission line.

The lower node of the left thermocouple is directly connected to ground and the lower node of the right thermocouple is at RF ground through bypass capacitor Cb. The DC voltages generated by the separate thermocouples add in series to form a higher DC output voltage. The principal advantage, however, of the two-thermocouple scheme is that both leads to the voltmeter are at RF ground; there is no need for an RF choke in the upper lead. If a choke were needed it would limit the frequency range of the sensor.

Bypass Capacitor

Thermocouple Chip

Input Blocking Capacitor

Housing

Sapphire Substrate

Figure 4-6. Sketch of the thermocouple assembly for the HP 8481A.

The thermocouple chip is attached to a transmission line deposited on a sapphire substrate as shown in Figure 4-6. A coplanar transmission line structure allows the typical 50 Ω line dimensions to taper down to the chip size, while still maintaining the same characteristic impedance in every cross-sectional plane. This structure contributes to the very low reflection coefficient of the HP 8480-series sensors, its biggest contribution, over the entire 100 kHz to 50 GHz frequency range.

The principal characteristic of a thermocouple sensor for high frequency power measurement is its sensitivity in microvolts output per milliwatt of RF power input. The sensitivity is equal to the product of two other parameters of the thermocouple, the thermoelectric power and the thermal resistance.

The thermoelectric power (not really a power but physics texts use that term) is the thermocouple output in microvolts per degree Celsius of temperature difference between the hot and cold junction. In the HP 8481A thermocouple sensor, the thermoelectric power is designed to be 250µV/° C. This is managed by controlling the density of n-type impurities in the silicon chip.

The thickness of the HP 8481A silicon chip was selected so the thermocouple has a thermal resistance 0.4° C/mW. Thus, the overall sensitivity of each thermocouple is 100 µV/mW. Two thermocouples in series, however, yield a sensitivity of only 160 µV/mW because of thermal coupling between the thermocouples; the cold junction of one thermocouple is heated somewhat by the resistor of the other thermocouple giving a somewhat smaller temperature gradient.

sensor.

21

10

Time

Figure 4-7. Zero drift of thermocouple and thermistor power sensors due to being grasped by a hand.

The thermoelectric voltage is almost constant with external temperature. It depends mainly on the temperature gradients and only slightly on the ambient temperature. Still, ambient temperature variations must be prevented from establishing gradients. The chip itself is only 0.8 mm long and is thermally short-circuited by the relatively massive sapphire sub-strate. The entire assembly is enclosed in a copper housing. Figure 4-7 depicts the superior thermal behavior of a thermocouple compared to a thermistor power sensor.

The thermoelectric output varies somewhat with temperature. At high powers, where the average thermocouple temperature is raised, the output voltage is larger than predicted by extrapolating data from low power levels. At a power level of 30 mW the output increases 3 percent, at 100 mW, it is about 10 percent higher. The circuitry in the HP power meters used with thermocouples compensates for this effect. Circuitry in the sensor itself compensates for changes in its ambient temperature.

The thermal path resistance limits the maximum power that can be dissipated. If the hot junction rises to 500° C, differential thermal expansion causes the chip to fracture. Thus, the HP 8481A is limited to 300 mW maximum average power.

The thermal resistance combines with the thermal capacity to form the thermal time constant of 120 microseconds. This means that the thermocouple voltage falls to within 37 percent of its final value 120 µs after the RF power is removed. Response time for measurements, however, is usually much longer because it is limited by noise and filtering considerations in the voltmeter circuitry.

The only significant aging mechanism is thermal aging of the tantalum nitride resistors. A group of devices were stress tested, producing the results of Figure 4-8. These curves predict that if a device is stressed at 300 mW for one year, the resistance should increase by about 3.5 percent. Nine days at a half watt would cause an increase in resistance of 2 percent. Aging accumulates. On the other hand, aging effects of the tantalum-nitride termination are compensated by use of the power calibration procedure, whereby a precision 1 mW, 50 MHz source is used to set a known level on the meter.

Figure 4-8. Results of step stress aging test show percent change in thermocouple resistance when left at various power levels continuously for various periods of time.

1.50 1.25

1.00

0.75

0.50

0.25 0

|  |  | % Change |  |  |  |
| --- | --- | --- | --- | --- | --- |
|  |  | 10 / 5 |  |  |  |
|  |  | // 2 'L//1 |  |  |  |
|  |  |  |  | 1 Y | ear |

0.01

0.1

1                                        10

Times (days)

100

1000

8

6

4

2

0

22

It is relatively easy to adapt this sensor design for other requirements. For example, changing each tantalum-nitride resistor to a value of 150 Ω yields a 75 Ω system. To enhance low frequency RF performance, larger blocking and bypass capacitors extend input frequencies down to 100 kHz. This usually compromises high frequency performance due to increased loss and parasitic reactance of the capacitors. The HP 8482A power sensor is designed for 100 kHz to 4.2 GHz operation, while the standard HP 8481A operates from 10 MHz to 18 GHz.

Power Meters for Thermocouple Sensors

Introduction of thermocouple sensor technology required design of a new power meter architecture which could take advantage of increased power sensitivity, yet be able to deal with the very low output voltages of the sensors. This led to a family of power meter instrumentation starting with the HP 435A analog power meter, to the HP 436A digital power meter.2,3,4,5 Some years later the dual-channel HP 438A was introduced, which provided for computation of power ratios of channels A and B as well as power differences of channels A and B. The most recent HP 437B power meter offered digital data manipulations with the ability to store and recall sensor calibration factor data for up to 10 different power sensors.

To understand the principles of the instrument architecture, a very brief description will be given for the first-introduced thermocouple meter, the HP 435A analog power meter. This will be followed by an introduction of HP’s newest power meters, HP E4418A (single channel) and HP E4419A (dual channel) power meters. They will be completely described in Chapter V, after a new wide-dynamic-range diode sensor is introduced.

Thermocouple sensor DC output is very low-level (approximately 160 nV for 1 microwatt applied power), so it is difficult to transmit in an ordinary flexible connection cable. This problem is multiplied if the user wants a long cable (25 feet and more) between the sensor and power meter. For this reason it was decided to include some low-level AC amplification circuitry in the power sensor, so only relatively high-level signals appear on the cable.

One practical way to handle such tiny DC voltages is to “chop” them to form a square wave, then amplify with an AC-coupled system. After appropriate amplification (some gain in the sensor, some in the meter), the signal is synchronously detected at the high-level AC. This produces a high-level DC signal which is then further processed to provide the measurement result. Figure 4-9 shows a simplified block diagram of the sensor/meter architecture.

Cabling considerations led to the decision to include the chopper and part of the first AC amplifier inside the power sensor. The chopper itself (Figure 4-10) uses FET switches that are in intimate thermal contact. This is essential to keep the two FET’s at exactly the same temperature to minimize drift. To eliminate undesired thermocouples only one metal, gold, is used throughout the entire DC path. All these contributions were necessary to help achieve the low drift already shown in Figure 4-7.

23

HP 8481A Power Sensor

HP 435A Power Meter

Figure 4-9. HP 435A/8481A architecture block diagram.

The chopping frequency of 220 Hz was chosen as a result of several factors. Factors that dictate a high chopping frequency include lower 1/f noise and a larger possible bandwidth, and thereby faster step response. Limiting the chopping to a low frequency is the fact that small transition spikes from chopping inevitably get included with the main signal. These spikes are at just the proper rate to be integrated by the synchronous detector and masquerade as valid signals. The fewer spikes per second, the smaller this masquerading signal. However, since the spikes are also present during the zero-setting operation, and remain the same value during the measurement of a signal, the spikes are essentially removed from the meter indication by zerosetting and cause no error. The spikes do, however, use up dynamic range of the amplifiers.

Figure 4-10. Simplified schematic of chopper amplifier.

200 fl

-vw-

II

O Thermocouple Output

DC Path (all gold)

To DC Amplifier

Pn

Multivibrator Input

One way to minimize noise while amplifying small signals is to limit the channel bandwidth. Since the noise generating mechanisms are broadband, limiting the amplifier bandwidth reduces the total noise power. The narrowest bandwidth is chosen for the weakest signals and the most sensitive range. As the power meter is switched to higher ranges, the bandwidth increases so that measurements can be made more rapidly. On the most sensitive range, the time constant is roughly 2 seconds, while on the higher ranges, the time constant is 0.1 seconds. A 2-second time constant corresponds to a 0 to 99 percent rise time of about 10 seconds.

24

Reference Oscillator

A frequent, sometimes well-directed criticism of thermocouple power measurements is that such measurements are open-loop, and thus thermistor power measurements are inherently more accurate because of their DC-substitution, closed-loop process. The bridge feedback of substituted DC power compensates for differences between thermistor mounts and for drift in the thermistor resistance-power characteristic without recalibration.

With thermocouples, where there is no direct power substitution, sensitivity differences between sensors or drift in the sensitivity due to aging or temperature can result in a different DC output voltage for the same RF power. Because there is no feedback to correct for different sensitivities, measurements with thermocouple sensors are said to be open-loop.

HP thermocouple power meters solve this limitation by incorporating a 50 MHz power-reference oscillator whose output power is controlled with great precision (±0.7 %). To verify the accuracy of the system, or adjust for a sensor of different sensitivity, the user connects the thermocouple sensor to the power reference output and, using a calibration adjustment, sets the meter to read 1.00 mW. By applying the 1 mW reference oscillator to the sensor’s input port just like the RF to be measured, the same capacitors, conductors and amplifier chain are used in the same way for measurement as for the reference calibration. This feature effectively transforms the system to a closed-loop, substitution-type system, and provides confidence in a traceability back to company and NIST standards.

HP EPM Series Power Meters

Figure 4-11. HP E4418A features many user-conveniences and a 90 dB dynamic measurement range.

The two-decade industry acceptance of HP thermocouple (and diode) sensor technology for RF power measurements has resulted in tens of thousands of units in the installed base around the world. Yet new technologies now allow for design of diode sensors with far larger dynamic range and new power meters with dramatically-expanded user features.

25

The HP E4418A (single channel) and E4419A (dual channel) power meters offer some significant user features:

• Menu-driven user interface, with softkeys for flexibility

• Large LCD display for ease of reading

• Sensor EEPROM which stores sensor calibration factors and other correction data (HP E series wide-dynamic-range CW sensors)

• Dedicated hardkeys for frequently-used functions

• Faster measurement speed, faster throughput

• Backward compatibility with all previous HP 8480-series sensors

• Form, fit, function replacement with HP 437B and 438A power meters (preserves automation software code)

• Built for wide-dynamic-range CW sensors from - 70 to +20 dBm.

Since the meters provide more powerful measurement capability when teamed with the HP E series wide-dynamic range diode sensors, the detailed description of the meters will be presented in Chapter V. This will first allow for a presentation of the technology for the new diode sensors with the range from - 70 to +20 dBm, and then the meters which take advantage of that increased capability.

Conclusions

Because of their inherent ability to sense power with true square-law characteristics, thermocouple sensors will always be best for handling signals with complex modulations or multiple tones. They always respond to the true average power of a signal, modulation, multiple signals and all. They are rugged and stable and reliable.

The large installed worldwide base of HP thermocouple and diode sensors and their compatible power meters argues that any new HP power meters be designed to be backward compatible with older sensors. All old HP 8480-series sensors will work with new HP EPM series meters, with the only limitation being the performance of the old sensors. The new HP E series sensors are not backwards-compatible with HP’s older meters due to a modified interface design, which allows for download of EEPROM-stored constants.

Thermocouple sensors are based on a stable technology that will be used to measure RF power in many applications for a long time in the future.

1. W.H. Jackson, “A Thin-Film Semiconductor Thermocouple for Microwave Power Measurements,” Hewlett-Packard Journal, Vol. 26, No. 1 (Sept., 1974).

2. A.P. Edwards, “Digital Power Meter Offers Improved Accuracy, Hands-Off Operation, Systems Capability,” Hewlett-Packard Journal, Vol. 27 No. 2 (Oct. 1975).

3. J.C. Lamy, “Microelectronics Enhance Thermocouple Power Measurements,” Hewlett-Packard Journal, Vol. 26, No. 1 (Sept., 1974).

4. “Power Meter-New Designs Add Accuracy and Convenience.” Microwaves, Vol. 13, No. 11 (Nov., 1974).

5. R.E. Pratt, “Very-Low-Level Microwave Power Measurements,” Hewlett-Packard Journal, Vol. 27, No. 2 (Oct., 1975).

26

V. Diode Sensors and Instrumentation

Rectifying diodes have long been used as detectors and for relative power measurements at microwave frequencies. The earliest diodes were used mostly for envelope detection and as nonlinear mixer components in superheterodyne receivers. For absolute power measurement, however, diode technology had been limited mainly to RF and low microwave frequencies.

High-frequency diodes began with point-contact technology which evolved from the galena crystal and cat-whisker types of early radio, and found application as early as 1904.1 Point-contact diodes were extremely fragile, not very repeatable, and subject to change with time. It was the low-barrier Schottky (LBS) diode technology which made it possible to construct diodes with metal-semiconductor junctions for microwave frequencies that were very rugged and consistent from diode to diode. These diodes, introduced as power sensors in 1974, were able to detect and measure power as low as - 70 dBm (100 pW) at frequencies up to 18 GHz.

This chapter will review the semiconductor principles as they apply to microwave detection, briefly review low-barrier Schottky technology and then describe the latest planar-doped-barrier (PDB) diode technology. It will describe how such devices are designed into power sensor configurations, and introduce a new CW-diode sensor with an impressive 90 dB dynamic power range using digital-detection-curve correction techniques. Signal and waveform effects for non-CW signals operating above the square-law range will also be examined.

The generic HP power meter family (HP 435, 436, 437, 438) was introduced in Chapter IV. This family has gained considerable importance because of the large installed base of HP meters and power sensors worldwide, all interoperable. In this chapter, two power meters will be described, which exploit the advantages of the new 90-dB-range sensors and offers major user-conveniences as well.

jiamps

Figure 5-1. The junction rectifying characteristic of a low-barrier Schottky diode, showing the small-signal, square-law characteristic around the origin.

Diode Detector Principles

Diodes convert high frequency energy to DC by way of their rectification properties, which arise from their nonlinear current-voltage (i-v) characteristic. It might seem that an ordinary silicon p-n junction diode would, when suitably packaged, be a sensitive RF detector. However, p-n junctions have limited bandwidth. In addition, the silicon p-n junction, without bias, has an extremely high impedance and will supply little detected power to a load. An RF signal would have to be quite large to drive the junction voltage up to 0.7 volts where significant current begins to flow. One alternative is to bias the diode to 0.7 volts, at which point it only takes a small RF signal to cause significant rectified current. This effort turns out to be fruitless, however, because the forward current bias gives rise to large amounts of noise and thermal drift. There is little, if any, improvement in the minimum power that can be metered.

Metal-semiconductor junctions, exemplified by point-contact technology, exhibit a low potential barrier across their junction, with a forward voltage of about 0.3 volts. They have superior RF and microwave performance, and were popular in earlier decades. Low-barrier Schottky diodes, which are metal-semiconductor junctions, succeeded point-contacts, and vastly improved the repeatability and reliability. Figure 5-1 shows a typical diode i-v characteristic of a low-barrier Schottky junction, expanded around the zero-axis to show the square-law region.

v

27

Mathematically, a detecting diode obeys the diode equation

av i = I      (e        –1) s

(5-1)

10v

1v

100mv

10mv

1mv

100|iv

10|iv

1|IV

100nv

-70 -60 -50 -40 -30 -20 -10      0 +10 +20 Input Power -dBm

-60 -50 -40 -30 -20 -10      0 +10 +20 Input Power -dBm

+2

0

-2

-4

-6

-8

-10

-12

14

Figure 5-2.

The diode detection

characteristic ranges

from square law

through a transition

region to linear

detection.

where a = q/nKT,

and i is the diode current, v is the net voltage across the diode, Is is the saturation current and is constant at a given temperature. K is Boltzmann’s constant, T is absolute temperature, q is the charge of an electron and n is a correction constant to fit experimental data (n equals approximately 1.1 for the devices used here for sensing power). The value

-of a    is typically a little under 40 (volts 1).

Equation (5-1) is often written as a power series to better analyze the rectifying action,

i = I      (av + s

(a v) 2 +      + .

2 !

v)3

3 !

(5-2)

It is the second and other even-order terms of this series which provide the rectification. For small signals, only the second-order term is significant so the diode is said to be operating in the square-law region. In that region, output i (and output v) is proportional to RF input voltage squared. When v is so high that the fourth and higher order terms become significant the diode response is no longer in the square law region. It then rectifies according to a quasi-square-law i-v region which is sometimes called the transition region. Above that range it moves into the linear detection region (output v proportional to input v).

For a typical packaged diode, the square-law detection region exists from the noise level up to approximately - 20 dBm. The transition region ranges from approximately - 20 to 0 dBm input power, while the linear detection region extends above approximately 0 dBm. Zero dBm RF input voltage is equivalent to approximately 220 mV (rms) in a 50 Ω system. For wide-dynamic-range power sensors, it is crucial to have a well-characterized expression of the transition and linear detection range.

Figure 5-2 shows a typical detection curve, starting near the noise level of - 70 dBm and extending up to +20 dBm. It is divided up into the square law, transition and linear regions. (Noise is assumed to be zero to show the square-law curve extends theoretically to infinitely small power.) Detection diodes can now be fabricated which exhibit transfer characteristics that are highly stable with time and temperature. Building on those features, data correction and compensation techniques can take advantage of the entire 90 dB of power detection range.

28

Figure 5-3. Circuit diagram of a source and a diode detector with matching resistor.

v0

The simplified circuit of Figure 5-3 represents an unbiased diode device for detecting low level RF signals. Detection occurs because the diode has a nonlinear i-v characteristic; the RF voltage across the diode is rectified and a DC output voltage results.

If the diode resistance for RF signals were matched to the generator source resistance, maximum RF power would be delivered to the diode. However, as long as the RF voltage is impressed across the diode, it will detect RF voltage efficiently. For reasons explained below, diode resistance for small RF signals is typically much larger than 50 Ω and a separate matching resistor is used to set the power sensor’s input termination impedance. Maximum power transfers to the diode when the diode resistance for small RF voltages is matched to the source resistance. The diode resistance at the origin, found by differentiating (5-1), is:

R o =

1

aI s

(5-3)

Resistance R      is a strong function of temperature which means the diode o

sensitivity and the reflection coefficient are also strong functions of temperature. To achieve less temperature dependence, R      is much larger than the

o

source resistance and a 50 Ω matching resistor serves as the primary termination of the generator. If R      of the diode of Figure 5-3 were made too

o

large, however, there would be poor power conversion from RF to DC; thus,

a larger R      decreases sensitivity. A compromise between good sensitivity o

to small signals and good temperature performance results from making I

s

about 10 microamps and R      between 1 to 2k Ω.

o

The desired value of saturation current, I , can be achieved by constructing

s

the diode of suitable materials that have a low potential barrier across the junction. Schottky metal-semiconductor junctions can be designed for such a low-barrier potential.

Using Diodes for Sensing Power

Precision semiconductor fabrication processes for silicon allowed the Schottky diodes to achieve excellent repeatability, and, because the junction area was larger, they were more rugged. HP’s first use of such a low-barrier Schottky diode (LBSD) for power sensing was introduced in 1975 as the HP 8484A power sensor.2 It achieved an exceptional power range from -70 dBm (100 pW) to -20 dBm (10 µW) from 10 MHz to 18 GHz.

As Gallium-Arsenide (GaAs) semiconductor material technology advanced in the 1980s, such devices exhibited superior performance over silicon in the microwave frequencies. A sophisticated diode fabrication process known as planar-doped-barrier (PDB) technology offered real advantages for power detection.3 It relied on a materials preparation process called molecular beam epitaxy for growing very thin epitaxial layers. Figure 5-4 shows the device cross sections of the two types of diode junctions, low-barrier Schottky (Figure 5-4 a) and planar-doped barrier

29

(Figure 5-4 b) diodes. The doping profile of the PDB device is n + —I—p + —I—n + , with intrinsic layers spaced between the n + and p + regions. The i/v characteristic has a high degree of symmetry which is related to the symmetry of the dopants.

The p + region is fabricated between the two intrinsic layers of unequal thickness. This asymmetry is necessary to give the PDB device the characteristics of a rectifying diode. An important feature of the PDB diode is that the device can be designed to give a junction capacitance, C o , that is both extremely small (20 fF or less) (fem to Farad) and nearly independent of the bias voltage. C o is also independent of the size of the metal contact pad.

As a result of the very stable C o vs bias voltage, the square-law characteristics of this device vs. frequency are significantly more constant than those of metal-to-semiconductor devices. Low capacitance coupled with low junction resistance allows excellent microwave signal performance since the low junction resistance lowers the RC time constant of the junction and raises the diode cutoff frequency.

A PDB diode is far less frequency-sensitive than a normal pn junction diode because of the intrinsic layer at the junction.4 In a pn junction diode the equivalent capacitance of the junction changes with power level, but in the planar-doped-barrier diode, junction capacitance is determined by the intrinsic layer, which remains almost constant as a function of power.

HP uses a specialized integrated circuit process which allows custom tailoring of the doping to control the height of the Schottky barrier. This controlled doping makes it practical to operate the detector diode in the current mode, while keeping the video resistance low enough to maintain high sensitivity.

The first power sensor application for PDB diodes was the HP 8481/85/87D-series power sensors, introduced in 1987.4 The HP 8481D functionally replaced the low-barrier-Schottky HP 8484A sensor. The new PDB sensor employed two diodes, fabricated on the same chip using MMIC (microwave monolithic integrated circuit) chip technology. The two diodes were deposited symmetrically about the center of a coplanar transmission line, and driven in a push-pull manner, for improved signal detection and cancellation of common-mode noise and thermal effects. This configuration features several advantages:

•    Thermoelectric voltages resulting from the joining of dissimilar metals, a serious problem below -60 dBm, are cancelled.

•    Measurement errors caused by even-order harmonics in the input signal are suppressed, due to the balanced configuration.

•    A signal-to-noise improvement of 1 to 2 dB is realized by having two diodes. The detected output signal is doubled in voltage (quadrupled in power) while the noise output is doubled in power, since the dominant noise sources are uncorrelated.

•    PDB devices also have higher resistance to electrostatic discharge, and are more rugged than Schottky’s.

•    Common-mode noise or interference riding on the ground plane is cancelled at the detector output. This is not RF noise, but metallic connection noises on the meter side.

PDB diode technology provides some 3000 times (35 dB) more efficient RF-to-DC conversion compared to the thermocouple sensors of Chapter IV.

30

(a) Small Signal

100

0

-50

-100

-50

-25

0 v (mv)

25

50

12.0 10.0 8.0 6.0 4.0 2.0

(b) Large Signal

-4.0 -3.5 -3.0-2.5 -2.0-1.5 -1.0 -0.5 0      0.5 1.0 v (volts)

Figure 5-5. The i-v characteristics of a PDB diode are shown for two different drive voltage regions. The

+ assymetry of the p

layer can be used to

modify the shape of

the i-v curve, while

not affecting some

of the other diode

parameters such as

C      and R . oo

Figure 5-5 shows two regions of the i-v characteristic of a typical PDB diode. Figure 5-5 (a) shows the small signal region, while Figure 5-5 (b) shows the larger signal characteristics to include the linear region as well as the breakdown region on the left.

They also provide accurate square-law performance from - 70 to - 20 dBm. Diode sensor technology excels in sensitivity, although realistically, thermocouple sensors maintain their one primary advantage as pure square-law detectors for the higher power ranges of - 30 to +20 dBm. Hence neither technology replaces the other, and the user’s measuring application determines the choice of sensors.

In detecting power levels of 100 picowatts (- 70 dBm) the diode detector output is about 50 nanovolts. This low signal level requires sophisticated amplifier and chopper circuit design to prevent leakage signals and thermocouple effects from dominating the desired signal. Earlier HP diode power sensors required additional size and weight to achieve controlled thermal isolation of the diode. The dual-diode configuration balances many of the temperature effects of those highly-sensitive circuits, and achieves superior drift characteristics in a smaller, lower-cost structure.

New Wide-Dynamic-Range CW-only Power Sensors

Digital signal processing and microwave semiconductor technology have now advanced to the point where dramatically-improved performance and capabilities are available for diode power sensing and metering. New diode power sensors are now capable of measuring over a dynamic power range of - 70 to +20 dBm, an impressive range of 90 dB. This permits the new sensors to be used for CW applications which previously required two separate

The new HP E4412A power sensor features a frequency range from 10 MHz to 18 GHz. HP E4413A power sensor operates to 26.5 GHz. Both provide the same - 70 to +20 dBm power range. A simplified schematic of the new sensors is shown in Figure 5-6. The front end construction is based on MMIC technology and combines the matching input pad, balanced PDB diodes, FET choppers, integrated RF filter capacitors, and the line-driving pre-amplifier. All of those components operate at such low signal levels that it was necessary to integrate them into a single thermal space on a surface-mount-technology PC board.

Figure 5-6. Simplified schematic for the new HP E series power sensors. The 90 dB power range is achieved using data stored in the individual sensor EEPROM which contains linearization, temperature compensation and calibration factor corrections.

Serial Data

50

sensors.

31

To achieve the expanded dynamic range of 90 dB, the sensor/meter architecture depends on a data compensation algorithm which is calibrated and stored in an individual EEPROM in each sensor. The data algorithm stores information of three parameters, input power level vs frequency vs temperature for the range 10 MHz to 18 or 26.5 GHz and - 70 to +20 dBm and 0 to 55° C.

At the time of sensor power-up, the power meter interrogates the attached sensor, using an industry-standard serial bus format, and in turn, receives the upload of sensor calibration data. An internal temperature sensor supplies the diode’s temperature data for the temperature-compensation algorithm in the power meter.

Since the calibration factor correction data will seldom be used manually, it is no longer listed on the sensor label of the HP E series sensors. The data is always uploaded into the power meter on power-up or when a new sensor is connected. The new sensors store cal factor tables for two different input power levels to improve accuracy of the correction routines. If the cal factor changes upon repair or recalibration, the new values are loaded into the sensor EEPROM. For system users who need the cal factor for program writing, the data is furnished on a calibration certificate.

A Versatile Power Meter to Exploit 90 dB Range Sensors

Figure 5-7. HP E4419A dual-channel power meter measures insertion loss of a 11.245 GHz waveguide bandpass filter, using its power ratio mode, plus a sweeper and power splitter. HP’s 90-dB dynamic-range sensors are ideal for such high-attenuation measurements.

Two power meters, HP E4418A (single channel) and HP E4419A (dual channel) take advantage of the sensor’s 90 dB power measuring range. More importantly, advances in digital signal processing (DSP) technology now provide significant increases in measurement speeds. Digital processing permits functional conveniences resulting in a dramatically more versatile power meter.

Figure 5-7 shows a front-panel view of the single-channel model. The main readout is an easy-to-read liquid crystal display with softkeys for utility and hardkeys for the main functional operation. Logical calibration and measurement sequences are grouped in the series of softkeys on the right of the display panel. The required 50 MHz power reference source is shown on the right top. The meter provides three measurement speeds, normal, x2 and fast mode. The speed is 20/40 readings per second for normal and x2 use, and 200 readings per second for automated system use. In the fast mode, the power range below - 50 dBm is not available because the noise filtering necessary in those ranges would slow the response to signal changes.

32

Figure 5-8. Simplified schematic of HP E4418A shows the takeover of digital signal processing (DSP) semiconductor architecture.

The basic meter architecture is based on DSP technology which improves performance by removing meter range switching and their delays (except for a single range-switching transition point). It also provides faster signal detection. The DSP module performs several other functions, synchronous detection (de-chopping), matches up the two analog-to-digital converter (ADC) channels, and does the programmable filtering. It provides a 32-bit digital number which is proportional to the detected diode voltage over a 50 dB power range.

The power meter uses the uploaded calibration data from each connected sensor to compensate for the three critical sensor parameters, power from - 70 to +20 dBm, frequency for its specified band, and operating temperature.

The calibration routine requiring connection to the 50 MHz power reference furnishes the traceable link for the sensor connected. The operator then keys in the frequency of the RF signal under test so that the meter corrects for the sensor calibration factor. Mismatch uncertainty must still be externally calculated because the reflection coefficient of the unknown power source is usually not available.

Figure 5-8 shows a simplified schematic of the HP E4418A meter. The pre-amplified sensor output signal receives some early amplification, followed by some signal conditioning and filtering. The signal is then split, with one path receiving amplification. Both low and high-level chopped signals are applied to a dual ADC. A serial output from the ADC takes the sampled signals to the digital signal processor which is controlled by the main microprocessor. A differential drive signal, synchronized to the ADC sampling clock, is output to the sensor for its chopping function.

AC Input

Feedback

Chop

-+—-

Temp Chop

Serial Data

The ADC provides a 20-bit data stream to the digital signal processor, which is under control of the main micro-processor. There is no range switching as in traditional power meters which maintain an analog signal path. Even the synchronous detection is performed by the ADC and DSP rather than use of a traditional synchronous detector.

Computation power permits the user to manipulate the basic measurement data to get desired units or format. Power reads out in watts or dBm, and inputs may be keyed in to compensate for attenuators or directional coupler losses to the unknown signal in front of the power sensor. Cabling losses can be compensated by entering the loss as a digital offset value.

33

For the two-channel power meter, either input power or both may be displayed. Or power ratio A/B or B/A might be useful for certain applications. For example, if the two power sensors are sampling forward and reverse power in a transmission using a dual directional coupler, these ratios would yield power reflection coefficient. The power difference, A- B or B- A, can be used for other applications. For example, using a dual directional coupler to sample forward and reverse power in a line, the power difference is a measure of net forward power being absorbed by a device under test. This is quite important in testing devices with very high reflections.

Power changes are displayed with the relative power function. And although the main display is all digital, a simple “peaking” display simulates an analog meter pointer and allows a user to adjust a unit under test for maximizing power output.

In system applications, the new single-channel power meter, when used with the wide-dynamic-range sensors can achieve 200 measurements per second. The programming code is also designed to be backward compatible with the previous HP 437B, (the E4419A is code compatible with the HP 438A). Of course, the new meter offers far more versatile programming functions too, to handle modern complex test procedures. But old software can be re-used to make programming projects more efficient.

When old sensors are utilized with the new meter, the calibration factor vs. frequency table printed on the label of the sensor must be keyed into the new power meters to take fullest advantage of the measurement accuracy. A table of frequencies vs. cal factor is displayed, and the routine prompted by the softkey display to ease editing.

Potential users of the new power meters will find that specification listings for this DSP-architecture meter without range switching will not follow traditional power meter range specs. Yet the meter meets the same range performance as the HP 43X- series meters.

Traceable Power Reference

All thermocouple and diode power sensors require a power reference to absolute power, traceable to the manufacturer or national standards. HP power meters accomplish this power traceability by use of a highly stable, internal 50 MHz power reference oscillator. When used together, the 50 MHz reference and the sensor calibration factor data supplied with each sensor yields the lowest measurement uncertainty. All HP sensors are supplied with calibration factor vs. frequency data. This includes both the value and uncertainty of each point.

For HP 435- 438-series meters, the 50 MHz sensitivity adjustment is made using the 1 mW, 50 MHz internal power reference. The calibration factor dial should then be adjusted to the proper value for the frequency of the signal under test. The calibration factor is marked on each HP 8480-series power sensor label or included data table. The HP 437B has a provision for user-storage of cal factor tables for up to 20 different sensors, allowing for quick change of sensors. Naturally, for each power measurement, the signal frequency must still be adjusted or keyed in.

34

The 1 mW reference power output is near the center of the dynamic range of thermocouple power sensors, but outside the range of the sensitive diode sensor series. A special 30 dB calibration attenuator, designed for excellent precision at 50 MHz, is supplied with each HP 8481D-series diode power sensor. When that attenuator is attached to the power reference output on the power meter, the emerging power is 1 µW (- 30 dBm). The attenuator design is such that a maximum error of 1 percent is added to the calibration step. Basic uncertainty of the reference output is factory adjusted to an accuracy of ±0.7% and is traceable to NIST. For 1 year, accuracy is ±1.2% worst case, ±0.9% rss.

Signal Waveform Effects on the Measurement Uncertainty of Diode Sensors

Along with the great increase in measurement flexibility of the E-series power sensor, comes several new applications guidelines. These must be understood and followed to obtain valid measurement results when dealing with complex and non-CW signals.

These guidelines distinguish between earlier diode sensors of the HP 8481D vintage and the E series CW-only diode sensors.

The power range from approximately - 20 to +20 dBm is above the square-law region, and one in which the EPM series power meters uses digital-diode-curve correction to provide accurate power measurement for pure CW signals all the way from - 70 to +20 dBm. The EPM meters and companion E series sensors provide fully specified performance over that entire dynamic range of - 70 to +20 dBm.

The following explanation reviews the effects of complex signals on existing HP 8481D-type diode sensors for non-CW or complex modulation signals.

Some examples of complex (non-CW) signals are as follows: 1) Pulsed RF such as radar or navigation formats, 2) Two-tone or multiple-tone signals such as those which might be present in a telecommunications channel with multiple sub-channels, 3) AM signals which have modulation frequencies higher than the bandwidth of the power meter detection filtering, (in the kHz range for the HP E4418A 4) Digital-phase-shift-keyed (PSK) modulations, 5) QAM modulated signals, 6) Pulse-burst formats.

Here is a summary of the measurement guidelines:

1) Using the HP 8481D type diode power sensors, any complex signal will yield highly-accurate measurement results as long as the peak power levels of the unknown signal are maintained below - 20 dBm. In addition, the lowest-frequency-component of any modulation frequency must be above approximately 500 Hz. Since the power range of the HP 8481D-type diode sensors are automatically restricted by HP power meters to a top level of - 20 dBm, the user need only see that no peak power levels go above - 20 dBm.

When peak power levels exceed approximately -20 dBm, accurate measurements can be accomplished by the simple expedient of attenuating the unknown signal through an external precise fixed or step attenuator, such that the complex signal peak power does not exceed -20 dBm. If pulse modulation frequencies are near the HP power meter chopping rate of 220 Hz or multiples thereof, some meter “beats” may be observed.

35

2) Using the new HP E series power sensors, CW signals will yield accurate results across their entire - 70 to +20 dBm dynamic range. One reason HP E series sensors may not be used for pulse power within their square-law range is that their output circuit filters are optimized for fast response to aid high data-rate automation.

3) For non-CW signals with average powers between - 20 and +20 dBm, use the HP thermocouple sensors for true average power sensing.

4) For complex signals and high rate modulation frequencies, such as pulsed radar or high-data-rate PSK modulations, the recommended measurement solution is the HP peak power analyzer, which is explained in detail in Chapter VIII.

It is quite easy to realize that thermal sensors such as the thermocouple are pure square law because they convert the unknown RF power to heat and detect that heat transfer. Conversely, it is less easy to understand how diode sensors can perform the square-law function without the heat transfer step in the middle. Diode detectors do deliver pure square-law performance in their lower power ranges below - 20 dBm, due to their mathematical detection transfer function, as described by the power series equation of (5-2).

A two-tone example might clarify the measurement example. Consider two CW signals, f1 and f2, of power level 0 dBm (1 mW) each, and separated by 1 MHz. In a 50 Ω system, each carrier would have voltage magnitudes of v1 = v2 = 0.223 volts. If the two-tone signal were measured by an HP 8481A thermocouple sensor, each carrier would convert the 1 mW into heat for a total of 2 mW.

Using a voltage vector analysis, these two-tone signals would be represented by a voltage minimum of zero and a voltage maximum of 0.446 volts, occurring at a frequency of 1 MHz. The problem then becomes evident when one realizes that 2 times voltage represents 4 times power. A shaped diode detector then interprets the 2V maximum as 4 times power, and averages it out to the wrong power reading.

Another example shows how subtle signal imperfections can cause errors. Consider a CW signal with a harmonic signal - 20 dBc (20 dB below the carrier amplitude or with a voltage equal to 10% of the carrier). Figure 5-9 shows a mathematical model of the increasing maximum error caused by a - 20 dBc harmonic signal, as the carrier power level ranges from - 30 to +20 dBm. While actual deviation from true power is a function of the phase difference between the carrier and harmonic, the error limits are shown to be as high as 0.9 dB. If the harmonic was measured in the true square-law region, a - 20 dBc harmonic represents only 1/100 th of the power of the carrier or 1% added power to the carrier.

It might also be observed that the design architecture of the PDB sensors utilizes a balanced, push-pull-diode configuration. This structure inherently rejects even-number harmonics of the RF input signal, therefore will provide 15 to 20 dB rejection of even-number harmonics above the square-law region.

36

Figure 5-9. Estimated error limits for diode detectors operated above square-law range, for CW signal with - 20 dBc harmonic.

0.5

-0.5

10 loq (1± —2-)2 ] 3                vi

20 log (1±      vv 21 )

w-20                      -10                        0

Power dBm

10                          20

Conclusion

The HP EPM series meters and E series sensors using detector-shaping compensation to deliver dynamic range above square law should only be used for CW signals.

Average power measurements on pulsed and complex modulation signals may be measured using HP thermocouple sensors and the HP EPM series power meters. HP 8480D-type diode sensors may be used below - 20 dBm.

Comprehensive characterization of signals with pulsed power and complex modulations should be made with a true peak power analyzer as reviewed in Chapter VIII.

1.    S.M. Sze, “Physics of Semiconductor Devices,” Second Edition, Wiley, (1981).

2. P.A. Szente, S. Adam, and R.B. Riley, “Low-Barrier Schottky-Diode Detectors,” Microwave Journal, Vol. 19 No. 2 (Feb., 1976).

3.    R.J. Malik, T.R. Aucoin and R.L. Ross, “Planar-Doped Barriers in GaAs Molecular Beam Epitaxy,” Electronics Letters, Vol 1G #22, (Oct., 1980).

4. A. A. Fraser, “A Planar-Doped-Barrier Detector for General Purpose Applications,” Microwave Journal, (May, 1987).

0

-1

37

VI. Measurement Uncertainty

In RF and microwave measurements there are many sources of measurement uncertainty. In power measurements, the largest errors are almost always caused by sensor and source mismatch. Mismatch uncertainties, which have several aspects, are somewhat complicated and are seldom completely understood or properly evaluated. This chapter begins with a description of uncertainties associated with mismatch. The concept of signal flowgraphs is introduced to aid in the visualization needed for understanding the mismatch process. Other sensor uncertainties such as effective efficiency and calibration factor are then considered. This is followed by an analysis of the various instrumentation uncertainties of the power meter. Finally, the chapter treats the combining of all errors for a total uncertainty number.

With regard to general treatment of uncertainty analysis, this chapter also briefly introduces the ISO Guide to the Expression of Uncertainty in Measurement. ISO is the International Organization for Standardization in Geneva, Switzerland. Recently, the National Conference of Standards Laboratories cooperated with the American National Standards Institute to adapt the ISO document to U.S. terminology. It is presently being introduced as the metrology industry document ANSI/NCSL Z540-2-1996, U.S. Guide to the Expression of Uncertainty in Measurement. HP has established a policy to transition to these new processes for determining measurement uncertainties.

Power Transfer, Generators and Loads

The goal of an absolute power measurement is to characterize the unknown power output from some source (generator). Sometimes the generator is an actual signal generator or oscillator where the power sensor can be attached directly to that generator. On other occasions, however, the “generator” is actually an equivalent generator. For example, if the power source is separated from the measurement point by such components as transmission lines, directional couplers, amplifiers, mixers, etc, then all those components may be considered as parts of the generator. The port that the power sensor connects to, would be considered the output port of the equivalent generator.

To analyze the effects of impedance mismatch, this chapter explains mathematical models which describe loads, including power sensors, and generators, which apply to the RF and microwave frequency ranges. The microwave descriptions begin by relating back to the equivalent low-frequency concepts for those familiar with those frequencies. Signal flowgraph concepts aid in analyzing power flow between an arbitrary generator and load. From that analysis, the terms mismatch loss and mismatch uncertainty are defined.

RF Circuit Descriptions

At low frequencies, methods for describing a generator include the Thevenin and Norton equivalent circuits. The Thevenin equivalent circuit of a generator, for example, has a voltage generator es in series with an impedance Zg as shown in Figure 6-1. For a generator, even if composed of many components, es is defined as the voltage across the output port when the load is an open circuit. Zg is defined as the impedance seen looking back into the generator when all the sources inside the generator are reduced to zero.

38

Figure 6-1. A Thevenin equivalent generator connected to an arbitrary load.

The power delivered by a generator to a load is a function of the load

impedance. If the load is a perfect open or short circuit, the power delivered

is zero. Analysis of Figure 6-1 would show that the power delivered to the

load is a maximum when load impedance Z      is the complex conjugate of the

,

generator impedance Z . This power level is called the “power available g

from a generator,” or “maximum available power,” or “available power.”

When Z      = (R      + jX      ) and Z      = (R      + jX ), are complex conjugates of each ,,,     ggg

other, their resistive parts are equal and their imaginary parts are identical

in magnitude but of opposite sign; thus R      = R      and X      = - X . Complex con,g                      ,g

jugate is written with an * so that Z      = Z * is the required relationship for

,g

maximum power transfer.

The Thevenin equivalent circuit is not very useful at microwave frequencies for a number of reasons. First, the open circuit voltage is difficult to measure because of fringing capacitance and the loading effect of a voltmeter probe. Further, the concept of voltage loses usefulness at microwave frequencies where it is desired to define the voltage between two points along a transmission path, separated by a significant fraction of a wavelength. Also, there are problems involved in discussing voltage in rectangular waveguide. As a result, the concept of power is much more frequently used than voltage for characterizing generators at RF and microwave frequencies.

The open circuit which defines the Thevenin equivalent voltage generator is useless for measuring power because the power dissipated in an open termination is always zero. The reference impedance used for characterizing RF generators is almost always 50 Ω. The reason for this is that 50 Ω is easy to realize over the entire frequency range of interest with a transmission line of 50 Ω characteristic impedance and with a reflection-less termination.

The standard symbol for characteristic impedance, Z , is also the standard

o

symbol for reference impedance. In some cases, for example, where, 75 Ω

transmission lines are used in systems with a 50 Ω reference impedance,

another symbol, such as Z , should be used for reference impedance. Z      will

ro

be used in this application note to mean reference impedance. A generator is

characterized, therefore, by the power it delivers to a reference load

Z ,= 50 Ω. In general, that power is not equal to the maximum available o

power from the generator; they are equal only if Z      = Z .

go

As frequencies exceed 300 MHz, the concept of impedance loses usefulness and is replaced by the concept of reflection coefficient. The impedance seen looking down a transmission line toward a mismatched load, varies continuously with the position along the line. The magnitude, and the phase of impedance are functions of line position. Reflection coefficient is well-behaved; it has a magnitude that is constant and a phase angle that varies linearly with distance from the load.

39

Reflection Coefficient

At microwave frequencies where power typically is delivered to a load by a transmission line that is many wavelengths long, it is very convenient to replace the impedance description of the load, involving voltage and current and their ratio (Ohm’s law), with a reflection coefficient description involving incident and reflected traveling waves, and their ratio. To characterize a passive load, Ohm’s law is replaced by:

b

,= G                                                                                                                                              (6-1)

a,

,

where a      is proportional to the voltage of the incident wave, b      is propor-,,

tional to the voltage of the reflected wave, and G        is defined to be the reflec-,

tion coefficient of the load. All three quantities are, in general, complex numbers and change with frequency. The quantities a      and b      are normalized1 in

,    ,

such a way that the following equations hold:

(6-2) ïa ï2 = P ,i

ï b,ï2 = Pr                                                                                                                                            (6-3)

where P is power incident on the load and P      is power reflected by it. ir

The net power dissipated by the load, P , is given by:

d

P      = P – P      = ïa ï2 – ïb ï2                                                                                  (6-4)

di              r,    ,

This power is the total power obtained from the source; it includes not only power converted to heat, but also power radiated to space, and power that leaks through accessory cables to other pieces of equipment.

Transmission line theory relates the reflection coefficient, G      , of a load to its

,

impedance, Z , as follows: ,

Z– Z ,o                                                                                                                                (6-5)

G= ,

Z+ Z ,o

where Zo is the characteristic impedance of the system. Further, the load

voltage, v , and load current i      are given by ,,

V,= Incident voltage + reflected voltage                                          (6-6)

= √ Zo (a,+ b,)

i        = Incident current – reflected current ,                                                                                                                                                                      (6-7)

1

VZ

(a      - b )

o

1. If the transmission line characteristic impedance is Z , the normalization factor is

o

√ Zo; that is, a, is obtained from the voltage of the incident wave by dividing by √ Zo. Similarly, b,is obtained from the voltage of the reflected wave by dividing by √ Zo.

40

since current in a traveling wave is obtained from the voltage by dividing

by Z . Solving for a      and b , results in: o                                              ,,

a      =

b      =

2\Z

2\Z

(v      + Z o i )

(v      - Z o i )

(6-8)

(6-9)

These equations are used in much of the literature to define a      and b      (see

,,

the reference by Kurakawa.)2 The aim here, however, is to introduce

a      and b      more intuitively. Although (6-8) and (6-9) appear complicated, the ,,

relationships to power (equations 6-2, 6-3, and 6-4) are very simple. The

Superposition Theorem, used extensively for network analysis,

applies to a      and b      ; the Superposition Theorem does not apply to power. ,    ,

Reflection coefficient G        is frequently expressed in terms of its magnitude ,

r        and phase f      . Thus r        gives the magnitude of b      with respect to a      and ,,,                                                                                  ,                                                        ,

f        gives the phase of b      with respect to a . ,                                                                  ,,

The most common methods of measuring reflection coefficient G        involve

,

observing a      and b      separately and then taking the ratio. Sometimes it is ,,

difficult to observe a      and b      separately, but it is possible to observe the ,,

interference pattern of the counter-travelling waves formed by a      and b      on

,,

a transmission line. This pattern is called the standing wave pattern. The

interference pattern has regions of maximum and of minimum signal

strength. The maximums are formed by constructive interference between

a      and b      and have amplitude | a      | + |b      |. The minimums are formed ,,                                                                                ,,

by destructive interference and have amplitude | a      | - | b      |. The ratio

,,

of the maximum to the minimum is called the standing-wave ratio (SWR,

sometimes referred to as voltage-standing-wave-ratio, VSWR) and can be

measured with a slotted line and moveable probe, or more commonly with

network analyzers. SWR is related to the magnitude of reflection coefficient

r        by: ,

SWR =

laj + IbJ          1 +|b/ aj            1 + p

-------1-------1-------1    = ---------1--------------1

aj–lbj          l–|b/ aj

1-P

(6-10)

Figure 6-2. Signal-flow graph for a load.

Signal Flowgraph Visualization

A popular method of visualizing the flow of power through a component or among various components is by means of a flow diagram called a signal flowgraph.1,2 This method of signal flow analysis was popularized in the mid-1960’s, at the time that network analyzers were introduced, as a means of describing wave travel in networks.

The signal-flow graph for a load (Figure 6-2) has two nodes, one to represent the incident wave a      and the other to represent the reflected wave b . ,,

They are connected by branch G        which shows how a      gets changed to

,,

become b . ,

1

1

41

Just as the Thevenin equivalent had two quantities for characterizing a generator, generator impedance and open circuit voltage, the microwave equivalent has two quantities for characterizing a microwave or RF generator, G        and b . The equation for a generator is: gs

b= b+ G      a g            s              gg

where:

b      is the wave emerging from the generator g

a      is the wave incident upon the generator from other components g

G        is the reflection coefficient looking back into the generator g

b    is the internally generated wave s

(6-11)

is related to Z      by g

Z– Z go

(6-12)

Z+ Z go

which is very similar to (6-5). b      is related to the power to a reference load

s

from the generator, P            , by gZo

P            = ïb ï gZo                  s

(6-13)

bsO

Figure 6-3. Signal-flow graph of a microwave generator.

Figure 6-4. The complete signal-flow graph of a generator connected to a load.

b    is related to the Thevenin voltage e      by: ss

√Z

b s =

e s

Z o + Z g

(6-14)

The signal-flow graph of a generator has two nodes representing the

incident wave a      and reflected wave b . The generator also has an gg

internal node b      that represents the ability of the generator to produce s

power. It contributes to output wave b      by means of a branch of value one.

g

The other component of b      is that portion of the incident wave a , that gg

is reflected off the generator.

Now that equivalent circuits for a load and generator have been covered,

the flow of power from the generator to the load may be analyzed. When the

load is connected to the generator, the emerging wave from the generator

becomes the incident wave to the load and the reflected wave from the load

becomes the incident wave to the generator. The complete signal-flow graph

(Figure 6-4) shows the identity of those waves by connecting node

b      to a      and node b      to a      with branches of value one. g,        ,g

Figure 6-4 shows the effect of mismatch or reflection. First, power from the generator is reflected by the load. That reflected power is re-reflected from the generator and combines with the power then being created by the generator, generating a new incident wave. The new incident wave reflects and the process continues on and on. It does converge, however, to the same result that will now be found by algebra.

The equation of the load (6-1) is rewritten with the identity of a      to b

g,

added as:

b , =r, a , = a g

(6-15)

I

g

42

The equation of the generator (6-11) is also rewritten with the identity of

a      to b      added as: ,g

b= b+ Ga= a                                                                                                          (6-16)

g              s              gg              ,

Equations (6-15) and (6-16) may be solved for a      and b      in terms of b , G

,,                                            s,

and G      : g

b s a      =                                                                                                                                                          (6-17)

,

1– G      G g,

bG s,

b      =                                                                                                                                                          (6-18)

, 1– GG

g,

From these solutions the incident and reflected powers can be calculated:

1 P = ïa ï2 = ïb ï2 i,     s                                                                                                                    (6-19)

ï1– G      G      ï2 g,

ïG        ï2 , P    = ïb ï2 = ïb ï2                                                                                                              (6-20)

r,    s

ï1– G      G      ï2 g,

Equation (6-19) yields the somewhat surprising fact that power flowing toward the load depends only on the load characteristics.

The power dissipated, P , is equal to the net power delivered by the gener-d

ator to the load, P gl

1– ïG        ï2 P      = P          = P– P= ïb ï2                                    ,        (6-21)

d              g,  i              r                s

ï1– G      G      ï2 g,

Two particular cases of equation (6-21) are of interest. First, if G        were

,

zero, that is if the load impedance were Z , equation (6-21) would give the

o

power delivered by the generator to a Z      load

o

P          ï                            = P              = ïb ï2                                                                          (6-22)

g,Z= Z                gZo                s

,o

This case is used to define b      as the generated wave of the source.

s

The second case of interest occurs when:

Gg= G,*                                                                                                                                            (6-23)

where * indicates the complex conjugate. Interpreting (6-23) means that the reflection coefficient looking toward the load from the generator is the complex conjugate of the reflection coefficient looking back toward the generator. It is also true that the impedances looking in the two directions are complex conjugates of each other. The generator is said to be “conju-gately matched.” If G        is somehow adjusted so that (6-23) holds, the ,

generator puts out its “maximum available power,” P        , which can be

av

expressed as:

ïb ï2 s Pav =                                  2                                                                                                              (6-24)

1– ïG        ï g

43

Comparing (6-22) and (6-24) shows that P          ³ P            .

av              gZo

Unfortunately, the term “match” is popularly used to describe both conditions, Z      = Z      and Z      = Z *. The use of the single word “match” should ,o                    ,g

be dropped in favor of “Z      match” to describe a load of zero reflection coeffi-o

cient, and in favor of “conjugate match” to describe the load that provides maximum power transfer.

Now the differences can be plainly seen. When a power sensor is attached to

a generator, the measured power which results is P        , of equation

g,

(6-21). But the proper power for characterizing the generator is P

gZo

of equation (6-22). The ratio of equations (6-22) to (6-21) is:

P gZo

P g,

ï1– G      G      ï2 g,

1– ïG        ï2

(6-25)

or, in dB:

dB = 10 log

P gZo

P g,

(6-26)

dB = 10 log ï1– G      G      ï2 – 10 log (1 – ïG        ï2) g,                                                          ,

This ratio (in dB) is called the “Z      mismatch loss.” It is quite possible that

o

(6-25) could yield a number less than one. Then (6-26) would yield a negative number of dB.

In that case more power would be transferred to the particular load being

used than to a Z      load, where the Z      mismatch loss is actually a gain. oo

An example of such a case occurs when the load and generator are conjugately matched.

A similar difference exists for the case of conjugate match; the measurement of P          from (6-21) differs from P          of (6-24). The ratio of those g,                                                                                      av

equations is:

P                                                    ï1– GGï2

av                                                g,

= Pg, (1 – ïG gï 2)(1 – ïG    ,ï 2)

(6-27)

or, in dB:

P av dB = 10 log

P g,

dB = 10 log ï1– G      G      ï2 – 10 log (1 – ïG        ï2) – 10 log (1 – ïG        ï2) ,g                                                        g                                                          ,

(6-28)

This ratio in dB is called the conjugate mismatch loss.

If G        and G        were completely known, there would be no difficulty. ,    g

The power meter reading of P          would be combined with the proper values

g,

of G        and G        in (6-25) or (6-27) to calculate P              or P        . The mismatch ,g                                                                                                          gZo                av

would be corrected and there would be no uncertainty.

44

Mismatch Uncertainty

G        and G        are seldom completely known. Only the magnitudes r        and r        are ,    g,    g

usually measured or specified. In these cases, the first term of the right side

of equations (6-26) and (6-28) cannot be exactly calculated because of the

lack of phase information, but the maximum and minimum values can be

found. The maximum and minimum values of 10 log|1 - G          G        | 2 are called

g,

“mismatch uncertainty limits” and are given the symbol M . The maximum

u

occurs when G        G        combines with “one” in phase to yield: ,g

M                      = 10 log (1 + r      r      )2

u max                                                  g,

(6-29)

This maximum limit will always be a positive number but it cannot be larger than 6 dB (this occurs when r        = r        = 1). The minimum value of the mis, g

match uncertainty occurs when G        G        combines with “one” exactly out of

,g

phase to yield:

M u min = 10 log (1 - pgp )2

(6-30)

The minimum limit will always be a negative number. It is also true that the magnitude of the minimum limit will be greater than the magnitude of the maximum limit, but usually by a very small amount.

Sometimes the mismatch uncertainty limits are given in percent deviation from “one” rather than in dB. In this case:

%M u = 100 [(1 ± pgp )2 - 1]

(6-31)

Mismatch uncertainty limits can be calculated by substituting the values

of r        and r        into equations (6-29), (6-30), and (6-31). ,g

The mismatch uncertainty limits can also be easily found by using the HP Reflectometer/ Mismatch Error Limits Calculator.6 This calculator also has slide rule scales for converting between SWR and r as well as other scales. Instructions and examples are printed on the calculator.

Figure 6-5. Reflectometer calculator and slide rule quickly determines limits of maximum and minimum mismatch uncertainty, if given the reflection coefficients of the generator and load. It is available from HP at nominal cost.

45

Modern engineering electronic calculators have a series of programs available especially suited for electrical engineering problems. One of the programs is intended for calculating mismatch uncertainty limits, either in terms of SWR or of r . Computer-aided engineering models often contain routines for transmission line calculations.

Mismatch Loss and Mismatch Gain

Traditionally, the transmission power loss due to signal reflection was termed mismatch loss. This was done in spite of the fact that occasionally the two reflection coefficient terms would align in a phase that produced a small “gain.” More recent usage finds the term mismatch gain more popular because it is a more inclusive term and can mean either gain (positive number) or loss (negative number). Similarly, it is more difficult to think of a negative mismatch loss as a gain. In this note, we use the terms interchangeably, with due consideration to the algebraic sign.

The second term on the right side of equation (6-26), - 10 log (1 - |G        | 2),

,

is called mismatch loss. It accounts for the power reflected from the load. In power measurements, mismatch loss is usually taken into account when correcting for the calibration factor of the sensor, to be covered below.

The conjugate mismatch loss of equation (6-28) can be calculated, if needed.

The uncertainty term is the same as the Z      mismatch loss uncertainty term

o

and the remaining terms are mismatch loss terms, one at the generator and one at the load. The term conjugate mismatch loss is not used much anymore. It was used when reflections were tuned out by adjusting for maximum power (corresponding to conjugate match). Now the various mismatch errors have been reduced to the point where the tedious tuning at each frequency is not worth the effort. In fact, modern techniques without tuning might possibly be more accurate because the tuners used to introduce their own errors that could not always be accounted for accurately.

Mismatch in power measurements generally causes the indicated power to be different from that absorbed by a reflection-less power sensor. The reflection from the power sensor is partially accounted for by the calibration factor of the sensor which is considered in the next chapter. The interaction of the sensor with the generator (the re-reflected waves) could be corrected only by knowledge of phase and amplitude of both reflection coefficients, G        and G      . If only the standing wave ratios or reflection ,    g

coefficient magnitudes r        and r        are known, then only the mismatch uncer-,    g

tainty limits can be calculated. The mismatch uncertainty is combined with all the other uncertainty terms later where an example for a typical measurement system is analyzed.

Other Sensor Uncertainties

After mismatch uncertainty, the second source of error is the imperfect efficiency of the power sensor. There are two parameters which define the design efficiency of a sensor, effective efficiency and calibration factor. Although HP now furnishes only calibration factor with its sensors, since both parameters are still available as measurement services for thermistor sensors from the National Institute of Standards and Technology, they will be reviewed here.

For a power sensor, the power input is the net power delivered to the sensor; it is the incident power minus the reflected power (P - P ). But not all

ir

that net input power is dissipated in the sensing element. Some might be radiated outside the transmission system or leaked into the instrumentation, some dissipated in the conducting walls of the structure, or in a

46

capacitor component of the sensor, or a number of other places that are not metered by the instrumentation. The metered power indicates only the power which is dissipated into the power sensing element itself.

For metering, the dissipated high frequency power must go through a conversion process to an equivalent DC or low frequency level. The DC or low

frequency equivalent is called P            , for substituted power. There are errors

sub

associated with the substitution process. In thermistor sensors, for example, errors result from the fact that the spatial distributions of current, power and resistance within the thermistor element are different for DC and RF power.

To accommodate both the usual parasitic losses as well as the DC or low frequency substitution problem mentioned, a special term, effective efficiency

h      , has been adopted for power sensors. Effective efficiency is defined by: e

P he=          sub                                                                                                                                        (6-32)

P g,

P          is the net power absorbed by the sensor during measurement. P              is g,                                                                                                                                                                                                                    sub

the substituted low frequency equivalent for the RF power being measured.

For thermistor sensors P              is the change in bias power required to bring sub

the thermistor back to the same resistance as before the application of RF

power. For thermocouple and diode sensors, P              is the amount of power

sub

from a reference power source, at a specified frequency, that yields the same

voltage to the metering circuits as P        . h        normally changes with frequen-g, e

cy, but changes with power level are usually negligible.

Effective efficiency is sometimes measured by the manufacturer when calibrating the sensor, and furnished in a calibration chart with the product. Sometimes the data is printed on the label of the sensor, or delineated with dots on a label plot of efficiency. It is expressed in percentage, and that factor is entered into the power meter by adjusting the analog dial to the appropriate number or entered digitally into digital power meters.

Calibration Factor

There is another more frequently used term that has been defined for power

measurements. It combines effective efficiency and mismatch loss and is

called the calibration factor K . K      is defined by:

bb

P sub Kb =                                                                                                                                                          (6-33)

P i

where P is the incident power to the sensor. The accurate measurement of i

calibration factor K      is quite involved and performed mainly by standards b

laboratories and manufacturers.

The definitions of K      and h        can be combined to yield be

P g, K= h                                = h        (1 – r2)                                                                          (6-34)

be                                      e                    ,

P i

where r        is the sensor reflection coefficient. The relationship on the right, ,

which is found by substituting for P and P          from equations (6-19) and

i                      g,

(6-21), shows that K      is a combination of effective efficiency and mis-b

match loss.

47

Most modern power meters have the ability to correct their meter reading

by setting a dial to the proper value of K . Then P is actually read off the

bi

meter. Values of K      for various frequencies are indicated on each Hewlett-b

Packard power sensor (except for the E series sensors which have the data

stored on EEPROM). When this feature is used, the indicated or metered

power P        is (using equation 6-19): m

PgZo                                      ïb ï2

s                                                                                      (6-35)

P=                        = P=

mi

Kb                                          ï1– GgG,ï2

But the desired quantity is usually not P to the sensor but P            , the

i                                                                gZo

power that would be dissipated in a Z      load. Since P              is by definition

o                                            gZo

| bs |2, the ratio of P              to the meter indication is: gZo

PgZo

= ï1– G      G      ï2                                                                                                                                  (6-36)

,g P m

The right side of (6-36) is the mismatch uncertainty. Since the use of K

b

corrects for efficiency and mismatch loss, only the mismatch uncertainty

remains. It should be pointed out that there is an additional, unavoidable

uncertainty associated with K . That uncertainty is due to inaccuracies in

b

the measurement of K      by the manufacturer, NIST or standards laborato-b

ries, and thus the uncertainty of K      is specified by the calibration supplier.

b

Power Meter Instrumentation Uncertainties

There are a number of uncertainties associated within the electronics of the

power meter. The effect of these errors is to create a difference between P

m

and P            /K . sub        b

Reference Oscillator Uncertainty

Open-loop power measurements, such as those that use thermocouples or semiconductor diode sensors, require a known source of power to verify and adjust for the sensitivity of the sensor. Many power meters, such as the HP 435A and 436A, have a stable power reference built in. No matter what power reference is used, if it deviates from the expected power output, the calibration adjustment is in error. The uncertainty in the power output from the reference oscillator is specified by the manufacturer. Thermistor power measurements, being closed-loop and having no need for a reference oscillator, are free of this error.

Reference Oscillator Mismatch Uncertainty

The reference oscillator has its own reflection coefficient at the operating frequency. This source reflection coefficient, together with that from the power sensor, creates its own mismatch uncertainty. Because the reference oscillator frequency is low, where the reflection coefficients are small, this uncertainty is small (approximately ±0.01 dB or ±0.2%).

Instrumentation Uncertainty

Instrumentation uncertainty is the combination of such factors as meter tracking errors, circuit nonlinearities, range-changing attenuator inaccuracy, and amplifier gain errors. The accumulated uncertainty is guaranteed by the instrument manufacturer to be within a certain limit.

There are other possible sources of uncertainty that are, by nature or design, so small as to be included within the instrumentation uncertainty.

48

An example of one such error is the thermoelectric voltage that may be introduced by temperature gradients within the electronic circuits and interconnecting cables. Proper design can minimize such effects by avoiding junctions of dissimilar metals at the most sensitive levels. Another example is the small uncertainty which might result from the operator’s interpolation of the meter indication.

±1 Count

On meters with digital output, there is an ambiguity in the least significant digit of ± one-half count. On some power meters, such as the HP 436A, this uncertainty is so small that it is absorbed in the instrumentation uncertainty. In some applications, such as relative power measurements or the ratio of two power measurements where most of the causes of instrumentation uncertainty do not affect the final result, this uncertainty is still applicable. In the case of relative power measurements, the uncertainty applies twice, once during the measurement of each power, for a total uncertainty of ± one count. One way of expressing the error is 1/P mant where P mant is the mantissa only of the meter indication. Another way is to find the relative power value of the least significant digit (lsd); the uncertainty is ±P lsd /P ind . This uncertainty can be reduced by using an external digital voltmeter of greater resolution.

Zero Set

In any power measurement, the meter must initially be set to “0” with no RF power applied to the sensor. Zero-setting is usually accomplished within the power meter by introducing an offset voltage that forces the meter to read zero, by either analog or digital means. The offset voltage is contaminated by several sources including sensor and circuit noise and setability of the zero set. The zero-set error is specified by the manufacturer, especially for the most sensitive range. On higher power ranges, error in zero setting is small in comparison to the signal being measured.

Zero Carryover

Most modern power meters, as a matter of convenience, have internal circuitry that eliminates the need to zero-set the power meter every time the power measurement range is changed. If the user zero-sets on the most sensitive range, they are then able to measure power on whatever range is of interest without re-zeroing. The circuitry that allows the zero-set to “carryover” to the other ranges may have slight offsets.

In principle, zero carryover uncertainty can be eliminated by zero-setting the power meter on the specific range of measurement. This practice is not recommended, however, for the HP 432A, 435A, 436A, 437B and 438A power meters. The EPM series meters do not have zero carryover. The zero carryover for these meters is typically much less than the data-sheet specification and the automatic zero-setting circuits operate more satisfactorily on the most sensitive range.

Noise

Noise is also known as short-term stability and it arises from sources within both the power sensor and circuitry. One cause of noise is the random motion of free electrons due to the finite temperature of the components. The power observation might be made at a time when this random fluctuation produces a maximum indication, or perhaps a minimum. Noise is specified as the change in meter indication over a short time interval (usually one minute) for a constant input power, constant temperature, and constant line voltage.

49

Drift

This is also called long-term stability, and is mostly sensor induced. It is the change in meter indication over a long time (usually one hour) for a constant input power, constant temperature, and constant line voltage. The manufacturer may state a required warm-up interval. In most cases the drift is actually a drift in the zero-setting. This means that for measurements on the upper ranges, drift contributes a very small amount to the total uncertainty. On the more sensitive ranges, drift can be reduced to a negligible level by zero-setting immediately prior to making a reading.

Power Linearity

Power measurement linearity is mostly a characteristic of the sensor. Deviation from perfect linearity usually occurs in the higher power range of the sensor. For thermocouple sensors, linearity is negligible except for the top power range of +10 to +20 dBm, where the deviation is specified at +2, - 4%.

For a typical HP 8481D-series diode sensor, the upper power range of - 30 to - 20 dBm exhibits a specified linearity deviation of ±1%.

With their much wider dynamic power range, the new HP E series sensors exhibit somewhat higher deviations from perfect linearity. It is mostly temperature-driven effect, and specifications are given for several ranges of temperature. For example, in the 25 ±5° C temperature range and the - 70 to - 10 dBm power range, the typical deviation from linearity is ±2% RSS.

Calculating Total Uncertainty

So far, only the individual errors have been discussed; now a total uncertainty must be found. This first description will use the traditional analysis for considering the individual uncertainties. It will be shortened to allow for later presentation of the recommended method for expressing uncertainties according to the ISO-based and U.S.-adapted considerations.

In some measurement applications, certain sources of error do not enter into the final uncertainty. An example of this is relative power measurement where the ratio of two power measurements is to be found. With proper procedure, the reference oscillator uncertainty affects the numerator and denominator in exactly the same way and therefore cancels out in the final result. In this same application, however, other errors might accumulate such as the ± half-count error.

50

Power Measurement Equation

The purpose of this section is to develop an equation that shows how a

power meter reading P        is related to the power a generator would deliver m

to a Z      load, P              (Figure 6-6). The equation will show how the individual o                        gzo

uncertainties contribute to the difference between P        and P            .

m                      gzo

Figure 6-6. Desired power output to be measured is P              , gZo

but measurement

results in the

reading P      . m

Starting from the generator in lower part of Figure 6-6, the first distinction is that the generator dissipates power P          in the power sensor instead

g,

of P              because of mismatch effects. That relationship, repeated from gZo

Chapter 5, is:

2

P gZo

ï1– G      G      ï

,g            P

g,

(6-37)

1 – ïG

The next distinction in Figure 6-6 is that the power sensor converts P        to

gl

the DC or low frequency equivalent, P            , for eventual metering. But this

sub

conversion is not perfect due to the fact that effective efficiency, h      , is less

e

than 100%. If P        is replaced by P            /h        from equation (6-32), then (6-37) gl                                                      sub      e

becomes:

P              = ï1– G      G      ï gZo                            ,g

Tle (1 -p2 )

P sub

(6-38)

The first factor on the right is the mismatch uncertainty term, M , dis-u

cussed previously. M      is also referred to as “gain due to mismatch”. The u

denominator of the second factor is the calibration factor K      from equation

b

(6-34). Now (6-38) can be written:

J_

1 PgZo = Mu                  P

K              sub b

(6-39)

The last distinguishing feature of Figure 6-6 is that the meter indication

P      , differs from P            . There are many possible sources of error in the m                                                  sub

power meter electronics that act like improper amplifier gain to the input

signal P            . These include uncertainty in range changing attenuators sub

and calibration-factor amplifiers, imperfections in the metering circuit and other sources totaled as instrumentation uncertainty. For open-loop power measurements this also includes those uncertainties associated with the calibration of amplifier gain with a power-reference oscillator. These errors are included in the symbol m for magnification.

1

51

| Table 1. Chart of Uncertainties For a Typical Absolute Power Measurement |  |  |  |  |
| --- | --- | --- | --- | --- |
| Measurement Conditions: P m = 50¦jW Full Scale (F.S.) = 100 pW p < 0.091 (SWR <1.2) pq<0.2 (SWR g<1.5) K[j = 93% ± 3% (worst case), ±1.5% (RSS) |  |  |  |  |
| Error | Description | Worst Case Values |  | RSS Component (∆X/X)2 |
|  |  | PgZo max | PgZo min |  |
| Mu | (1 ± PgP )2 | 1.0367 | 0.9639 | (0.0367)2 |
| K Uncertainty b | ±3% (w.c), ±1.5% (RSS) | 0.97 | 1.03 | (0.015)2 |
| Components of m: Ref. Osc. Unc. Ref. Osc. Mu Instrumentation Total m | ±1.2% SWRg = 1.05, SWR^ = 1.1 ±0.5% of F.S. | 0.988 0.998 0.99 0.9762 | 1.012 1.002 1.01 1.0242 | (0.012)2 (0.002)2 (1.01)2 |
| Components of t: Zero Set Zero Carryover Noise Total t | ±0.5% F.S. (low range) ±0.2% of F.S ±0.025 nW | -0.05 nW -0.2 nW -0.025 nW -0.275, nW | +0.05 nW +0.2 nW +0.025 nW +0.275 nW | (0.001)2 (0.004)2 (0.0005)2 |
| Expressions of total uncertainty: P gZo max P gZo min ∆PgZo ∆PgZo /Pm dB | equation (6-44) equation (6-45) | 55.0421 nW 5.0421 nW +10.08% 0.4171 dB | 45.4344 nW —4.5656 uW -9.13% -0.4159 dB | [0.001837) 1/2 = ±4.3% +0.1823 dB -0.1903 dB |

There are other uncertainties associated with the electronics that cause

deviation between P        and P            . When P              is zero, then P        should be m                      sub                              sub                                                  m

zero. Improper zero-setting, zero carryover, drift and noise are likely contributors to P        not being zero. The meter reading is offset or translated m

from mP              by a total amount t. A general linear equation gives P        in sub                                                                                                                                                                                    m

terms of P            : sub

P= mP+ t                                                                                                                    (6-40)

m                    sub

Substituting (6-40) into (6-39) gives the power measurement equation:

M      (P        – t) u          m

P gZo

Kbm

(6-41)

In the ideal measurement situation, M      has the value of one, the mK

u                                                                                                  b

product is one, and t is zero. Under ideal conditions, meter reading P

m

gives the proper value of P            .

gZo

52

Worst-Case Uncertainty

One method of combining uncertainties for power measurements in a worst-case manner is to add them linearly. This situation occurs if all the possible sources of error were at their extreme values and in such a direction as to add together constructively and therefore achieve the maximum possible

deviation between P      , and P            . Table 1 is a chart of the various error m                        gZo

terms for the power measurement of Figure 6-6. The measurement conditions listed at the top of Table 1 are taken as an example. The conditions and uncertainties listed are typical and the calculations are for illustration only. The calculations do not indicate what is possible using the most accurate technique. The description of most of the errors is from a manufacturer’s data sheet. Calculations are carried out to four decimal places because of calculation difficulties with several numbers of almost the same size.

Instrumentation uncertainty, i, is frequently specified in percent of full scale

(full scale = P      ). The contribution to magnification uncertainty is: fs

(1 + i) P fs

mi =                                                                                                                                                          (6-42)

P m

The several uncertainties that contribute to the total magnification uncertainty, m, combine like the gain of amplifiers in cascade. The minimum possible value of m occurs when each of the contributions to m is a minimum. The minimum value of m (0.9762) is the product of the individual factors (0.988 * 0.998 * 0.99).

The factors that contribute to the total offset uncertainty, t, combine like voltage generators in series; that is, they add. Once t is found, the contribution in dB is calculated from:

t tdB = 10 log (1 ±)                                                                                                              (6-43)

P m

The maximum possible value P            , using (6-41) and substituting the values

gZo

of Table 1, is

M                      (P– t              )

u max          m            min

PgZo max=                                                                                                                                    (6-44)

K                  m

b min          min

1.0367 (50 µW + 0.275 µW) =

(0.97) (0.9762)

= 55.0421 µW = 1.1008 P m

In (6-44), the deviation of Kbm from the ideal value of one is used to calculate P              max. In the same way, the minimum value of P              is: gZo                                                                                                                                                              gZo

M                      (P– t                )

u min          m            max

PgZo min=                                                                                                                                    (6-45)

K                  m

b max          max

0.9639 (50 µW – 0.275 µW)

(1.03) (1.0242)

= 45.4344 µW = 0.9087

53

The uncertainty in P                may be stated in several other ways: gZo

(1) As an absolute differential in power:

+5.0421

DP              = P                    –P        =        µW

gZo              gZo                  m

max                                –4.5656

min

(2) As a fractional deviation:

(6-46)

+5.0421 DP

gZo              –4.5656                +0.1008

P m

50

-0.0913

(6-47)

Figure 6-7. Graph of individual contributions to the total worst-case uncertainty.

| +0.5 |  |  |  |  |  |  |  |  | AP Zo |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  |  |  | \ ^« |  |  |
|  | Ref Osc j |  |  |  |  |  |  |  |  |
| +0.3 | Instrument] |  |  |  |  |  | Magnification |  |  |
|  |  |  |  | - | ^ |  |  |  |  |
| +0.2 |  |  |  |  | Calibr |  | ation Facto | ■ |  |
|  |  | ■• |  | - |  |  |  |  |  |
|  |  |  | ■ IV | ismatch |  |  |  |  |  |
| 0 | Pm | » |  |  |  |  |  |  |  |

(3) As a percent of the meter reading:

100 x

PgZo            +10.08

P m

-9.13

%

(6-48)

(4) As dB deviation from the meter reading:

dB = 10 log

1 .1 0 08 (0 .9 0 87)

+0.4171 -0.4159

dB

(6-49)

An advantage to this last method of expressing uncertainty is that this number can also be found by summing the individual error factors expressed in dB.

Figure 6-7 is a graph of contributions to worst-case uncertainty shows that mismatch uncertainty is the largest single component of total uncertainty. This is typical of most power measurements. Magnification and offset uncertainties, the easiest to evaluate from specifications and often the only uncertainties evaluated, contribute less than one-third of the total uncertainty.

RSS Uncertainty

The worst-case uncertainty is a very conservative approach. A more realistic method of combining uncertainties is the root-sum-of-the-squares (RSS) method. The RSS uncertainty is based on the fact that most of the errors of power measurement, although systematic and not random, are independent of each other. Since they are independent, it is reasonable to combine the individual uncertainties in an RSS manner.

54

Finding the RSS uncertainty requires that each individual uncertainty be expressed in fractional form. The RSS uncertainty for the power measurement equation (6-41) is:

D PPggZZoo = [(D MMuu)2 + (D KKbb)2 + ( D mm )2 + (P D mt )2]1/2

(6-50)

Each of the factors of (6-50), if not known directly, is also found by taking the RSS of its several components. Thus:

[( Dmm1)2 + (D mm 2)2 + • • •]1/2                                  (6-51)

m              [( D m1    2        ( D m22

D m= uf) + uf

Where m , m , and so forth are the reference oscillator uncertainty, the 12

instrumentation uncertainty, and so forth of Table 1.

The extreme right hand column of Table 1 shows the components used to find the total RSS uncertainty. The result is ±4.3%, which is much less than the worst case uncertainty of +10.1%, - 9.1%. One characteristic of the RSS method is that the final result is always larger than the largest single component of uncertainty.

New Method of Combining Power Meter Uncertainties

This section will describe a new method of combining uncertainties for HP power measurements. It is extended from the traditional uncertainty model, to follow a new guideline published by the American National Standards Institute and the National Conference of Standards Laboratories. The methods described in this document, ANSI/NCSL 540Z-2-1996, U.S. Guide to the Expression of Uncertainty in Measurement, are now being implemented for many metrology applications in industry and government.4

Generally, the impact of the ANSI\NCSL 540Z-2 guide is to inject a little more rigor and standardization into the metrology analysis. Traditionally, an uncertainty was viewed as having two components, namely, a random component and a systematic component. Random uncertainty presumably arises from unpredictable or stochastic temporal and spatial variations of influence quantities. Systematic uncertainty arises from a recognized effect or an influence which can be quantified.

The 540Z-2 guide groups uncertainty components into two categories based on their method of evaluation, Type “A” and Type “B.” These categories apply to uncertainty, and are not substitutes for the words “random” and “systematic.” Some systematic effects may be obtained by a Type A evaluation while in other cases by a Type B evaluation. Both types of evaluation are based on probability distributions. The uncertainty components resulting from either type are quantified by variances or standard deviations.

Briefly, the estimated variance characterizing an uncertainty component obtained from a Type A evaluation is calculated from a series of repeated measurements and is the familiar statistically estimated variances2 . Since standards laboratories regularly maintain measured variables data on their standards, such data would usually conform to the Type A definition.

55

For an uncertainty component obtained from a Type B evaluation, the estimated variance u2 is evaluated using available knowledge. Type B evaluation is obtained from an assumed probability density function based on the belief that an event will occur, and is often called subjective probability, and is usually based on a pool of comparatively reliable information. Others might call it “measurement experience.” Published data sheet specifications from a manufacturer would commonly fit the Type B definition.

Power Measurement Model for ISO Process

Beginning with the measurement equation of (6-41),

M      (P– t)

um                                                                                                                        (6-52)

P= gZo

Km b

The determination of m is through the calibration process. During calibration, P              is set to the known power, P          . Substituting P            for P              and gZo                                                                                                cal                                                    cal                  gZo

rearranging equation (6-52), equation for m is:

M          (P– t) m =                    uc          mc                                                                                                                (6-53)

KP c    cal

where:

m            = power meter gain term

M          = gain due to the mismatch between the sensor and the internal uc

calibration power source

P              = power level indicated by the power meter during calibration mc

t                = power meter zero offset

K            = power sensor calibration factor at the calibration frequency c

P            = power delivered to a Z      load by the power meter calibration output cal                                                                                  o

In equations (6-52) and (6-53), t represents the power meter zero offset.

In AN64-1, t is described as the sum of the zero set value, Z , zero carry-s

over, Z , Noise, N, and Drift, D. However, assuming the zero procedure c

occurs just prior to calibration, D is zero during calibration, whereas D

is non-zero during power meter measurements. To allow t to represent the

same quantity in the equation for P              and m, the equation for t is defined

gZo

as:

t = Z+ Z+ N                                                                                                                        (6-54)

sc

where,

Z      = power meter zero set value s

Z      = power meter zero carryover value c

N = power meter noise

and the equation for P              is redefined as gZo

M      (P        – (t+D)) u          m PgZo=                                                                                                                                                    (6-55)

Km b

where D = power meter drift.

56

Equation (6-55) is the measurement equation for a power meter measurement. There are eleven input quantities which ultimately determine the

estimated value of P            . These are M , P      , D, K      from equation (6-55); gZo                                              um                    b

P              K    and P            from equation ,        c                      cal

Z , Z , N from equation (6-54); and M sc                                                                                                        uc

(6-53). It is possible to combine equations in order to represent P              in

gZo

terms of the eleven defined input quantities. This is a relatively complicated derivation, but the result is the uncertainty in terms of the eleven quantities:

[u (mu)

u (P gZo ) = P gZo              M 2

u2 (P      ) m

u2 (D)

(P      – (t + D))2              (P      – (t + D))2 mm

u      (K b ) K2 b

u2 (P          )              u2 (K ) mc                                c

+

(P          – t)2                K2

mc                                      c

u2 (M        ) uc

M        2 uc

u2 (P          ) cal

P          2 cal

(P          – (t+D))2                  (P              – t)2                K P          m(P          – (t+D))

m                                                          mc                                    c    cal              m

)    u2(Zs) + u2(Zc) + u2(N)) ]

(6-56)

Solving with some nominal values of several input quantities simplifies equation (6-56),

|  | M u = 1 |  |
| --- | --- | --- |
|  | M uc = 1 |  |
|  | P mc = P cal |  |
|  | £-s = 0 |  |
|  | Z c = 0 |  |
|  | N = 0 |  |
|  | D = 0 |  |
|  | t = 0 |  |
|  | m = 1/K c |  |
| u2(P gZo ) | 2 u2 (P ) u(M)+ m u+ | u2 (D) + u2 (Kb) + u2 (M ) + u2 (Pmc) uc |
| P 2 gZo | P 2 m | P 2 K2 P 2 mb mc |
|  | u2 (K) u2 (P ) cl ++ K2 P2 cl | u Pcal) ( 1 1 ) 2 + 2+P P (u (Z s ) + u2(Z |

(6-57)

+

+

+

+

2

(

1

1

+

Table 2 summarizes the various uncertainties shown in 6-57.

57

| Table 2. Standard Uncertainties |  |
| --- | --- |
| Standard Uncertainty | Source |
| u(M u ) | Mismatch gain uncertainty between the sensor and the generator. The standard uncertainty is |
|  | dependent upon the reflection coefficients of the sensor and the generator. Refer to the mis- |
|  | match model. Reflection coefficients may have different distributions as shown in Figure 6-9. |
| u(M uc ) | Mismatch gain uncertainty between the sensor and the calibrator output of the power meter. |
|  | The standard uncertainty is dependent upon the reflection coefficients of the sensor and the |
|  | calibrator output. Refer to the mismatch model. Note: The calibrator output reflection coefficient |
|  | is not a specified parameter of the HP E4418A power meter. AN64-1 suggests r = 0.024. g |
| u(P m ) | Power meter instrumentation uncertainty. |
| u(P mc ) | Power meter instrumentation uncertainty (during calibration) |
| u(D) | Power meter drift uncertainty. |
| u(K b ) | Sensor calibration factor uncertainty. Typically, the value of the uncertainty is reported along |
|  | with the calibration factor by the calibration laboratory or the manufacturer. |
| u(K c ) | Sensor calibration factor uncertainty at the frequency of the power meter calibrator output. |
|  | If the sensor is calibrated relative to the associated calibrator output frequency, K = 1 and c |
|  | u(K c ) = 0. |
| u(Pl) | Power sensor linearity which is related to power range. Generally negligible on lower ranges |
|  | but has higher uncertainty at high power levels. |
| u(P cal ) | Calibrator output power level uncertainty. |
| u(Z s ) | Power meter zero set uncertainty. |
| u(Z c ) | Power meter zero carryover uncertainty. |
| u(N) | Power meter and sensor noise uncertainty. |

Standard Uncertainty of the Mismatch Model

The standard uncertainty of the mismatch expression, u(M ), assuming no

u

knowledge of the phase, depends upon the statistical distribution that best

represents the moduli of G        and G      .

g                      ,

Combining equation (6-21) and (6-22), the power dissipated in a load when

G        is not 0 is: ,

P= P d              gZo

ï1– G        G      ï2 g      ,

ï1– G      ï2

(6-58)

The denominator in (6-58) is known as mismatch loss. And the numerator represents the mismatch uncertainty:

M u

1– G      G      ï g,

(6-59)

M      is the gain or loss due to multiple reflections between the generator u

and the load. If both the moduli and phase angles of G                        G        are known,

g and      ,

M      can be precisely determined from equation (6-59). Generally, an esti-u

mate of the moduli exists, but the phase angles of G        and G        are not

g,

known.

58

Consider two cases:

Case (a): Uniform T, uniform phase distribution.

See Figure 6-8 (a). The moduli of      g and T      are each less than a specified value;

g and r      each lie within a circle of radius, T. Assuming      g and T      have equal probability of lying anywhere within the circle, the standard uncertainty of M u is: (This results in uniform density.)

Figure 6-8. When the reflection coefficients of the generator and load are not known, the user may estimate probabilities of the mismatch uncertainty according to these two cases; (a) both T lie inside the circle with uniform density. (b) both T lie on the circle with uniform phase density.

u(Mu) = 1      2 x maximum | rg | x maximum | T    |

Case (b): Constant p, uniform phase distribution.

See Figure 6-8 (b). An estimate of the moduli of      g and T      are known;      g and T each lie on a circle of radius T. Assuming      g and T      have equal probability of lying anywhere on the circle, (equal probability of any phase), the standard uncertainty of M u is:9

u(M u )

2 x r„ x r

Example of Calculation of Uncertainty Using ISO Model

Recognizing that each uncertainty calculation must meet a particular measuring requirement, the user will need to structure their calculations for appropriate conditions. This following measurement situation reflects some assumed and stated conditions for each of the parameters. The power meter is assumed to be the HP E4418A power meter, and the power sensor is assumed to be the HP E4412A power sensor.

Measurement conditions for calculation: Unknown CW power, 2 GHz, 50 microwatt level (-13 dBm).

Calculation comments for each parameter:

u(M u )

Uncertainty of mismatch gain between sensor and generator at 2 GHz. Use case (a) and assume generator reflection coefficient specification (from data sheet) is |      g | = 0.1 (uniform density distribution). Assume the HP E4412A sensor cal data shows a measured value of |T    | = 0.1 (uniform distribution of phase.)

Use mismatch gain equation of M u = |1 ±      g Y      | 2

Since each T has a different distribution, use a Monte Carlo

simulation.

u(Mu) M„

=      0.1 x 0.1 x

1

-{T

■= 0.7% (1-sigma)

u(M,

Uncertainty of mismatch gain between sensor and 50 MHz calibrator source. Use case (a) and assume source reflection coefficient specification (from data sheet) is |      g | = 0.024. HP E4412A sensor cal data shows a measured value of | T    | = 0.1 (uniform distribution of phase.) Use mismatch gain equation as above.

u(Mur)            ......                  1

M„

= 0.024 x 0.1 x

VX

= 0.17% (1-sigma)

59

u(P m )                HP E4418A power meter instrumentation uncertainty is

specified at ± 0.5%. (rectangular distribution) Use v 3 for divisor.

u(PJ                0.005

—                    = —1=—= 0.3% (1-sigma)

u(P mc )              HP E4418A power meter uncertainty during calibration. Specified at ±0.5%. (rectangular distribution)

u(Pmr)                0.005

= —1==—= 0.3% (1-sigma)

^mc                      ' 3

u(D)                      HP E4418A power meter drift uncertainty. Due to sensor drift.

Assume constant temperature, measurement taken one hour after calibration. From data sheet E series sensors are ±15pW. (rectangular distribution)

U(D)                15 X    10"12                  1                  n                                            ■

p      = soxio-6 x vr= °-0017% (1-Sl«ma)

u(K b )                  HP E4412A power sensor calibration factor uncertainty at

2 GHz. From the calibration certificate, spec is ±2%. (Gaussian distribution, 2-sigma)

u(Kk)                  0.02

—                    = —t\— =      1% (1-sigma)

Kb                          2

(For power levels from 0 to +20 dBm, an additional term u(K b ) should be RSSed (as a rectangular distribution) to account for high-power uK b uncertainty. For HP E-series sensors only.)

u(K c )                    HP E4412A power sensor cal factor uncertainty at 50 MHz is

assumed to be 0 since it is referred to the internal calibration

u(Kc) Kc

=0

u(P l )                      HP E4412A power sensor linearity uncertainty. For the

100 µW assumed range, this is specified for 25 ±5° C as ±4%. Assume cal lab temperatures within 5° C. Assume Gaussian distribution, 2-sigma.

u(Pi)                  0.04

—^,              = —t\— = 2% (1-sigma) Pi                            2

u(P cal ) 50 MHz calibrator power reference output uncertainty is specified at 0.9%, RSS, for 1 year. Guassian distribution, 2-sigma.

u(Pral)              0.009

= —-— = 0.45% (1-sigma)

u(Z s )                    HP E4418A power meter zero set uncertainty is specified at

±50pW. (rectangular distribution)

u (Zs) (—--------——)= I----------—--------— 1 x —j=---- = 0.00005% (1-sigma)

\Pm              Peal/        V50 x 10-              10-    /                V3

source.

60

u(Z )                    HP E4418A power meter zero carryover is included in the

c

overall instrument uncertainty specification, since there are no ranges as such in this meter. For other power meters this would need to be considered.

u (Zc) I ■

1

■pb)-0

u(N)                      HP E4418A power meter noise uncertainty is ±70pW and

negligible at the 50 mW power level.

u (N) I

1

1      W              1________1_\          70 xl012

Peal/        \50 xl0-6            10-3/X        V~3~

0.00007% (1-sigma)

Using the above comments, Table 3 summarizes the various uncertainty factors. Each factor is normalized to a 1 sigma value. In the case of a data sheet specification, the divisor factor used to convert to 1 sigma is square root of 3. These sigma values are added in RSS fashion, and then multiplied with the coverage factor. The coverage factor is a guard band number, typically 2 is used, but experience and knowledge of the measurement process allows for the user to establish any other value.

| Table 3. Worksheet for Uncertainties Calculation Using ISO Process |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Symbol | Source of Uncertainty | Value ±% | Probability Distribution | Divisor | D (K ) x |
| M u | Mismatch gain between generator and sensor | r = 0.1 g = 0.1 | ¦G ¦ - uniform density g ¦G ¦ - uniform phase s | (1) | 0.7% |
| M uc | Mismatch gain between calibration source and sensor | T ¦ = 0.024 gg ¦ = 0.1 | ¦G ¦ - uniform density g ¦G ¦ - uniform phase s | (1) | 0.17% |
| P m | Power meter instrumentation | 0.5% | rectangular | V3 | 0.29% |
| P mc | Power meter instrumentation during calibration | 0.5% | rectangular | v3 | 0.29% |
| D | Power meter drift | ±15pW | rectangular | v3 | 0.0017% |
| b | Sensor calibration factor | 2.0% | Gaussian | 2 | 1.0% |
| K c | Sensor calibration factor at 50 MHz | 0 | rectangular | — | 0 |
| P l | Power sensor linearity | 4.0% | Gaussian | 2 | 2.0% |
| P cal | Calibrator output power | 0.9% | Gaussian | 2 | 0.45% |
| Z s | Power meter zero set | ±50pW | rectangular | V3 | 0.00005% |
| Z c | Power meter zero carryover | 0 | rectangular | v3 | 0 |
| N | Power meter and sensor noise | ±70pW | rectangular | v3 | 0.00007% |
|  | Combined uncertainty—RSSed |  |  |  | 2.43% |
|  | Expanded uncertainty |  | Coverage factor K = 2 |  | 4.86% |

(1) Monte Carlo simulation

61

1.    N.J. Kuhn, “Simplified Signal Flow Graph Analysis,” Microwave Journal, Vol 6, No 10, Nov. 1963.

2. K. Kurakawa, “Power Waves and the Scattering Matrix,” IEEE Trans. on Microwave Theory and Techniques, Vol. 13, No. 2, Mar 1965.

3. “ISO Guide to the Expression of Uncertainty in Measurement,” International Organization for Standardization, Geneva, Switzerland, ISBN 92-67-10188-9, 1995.

4.    ANSI/NCSL Z540-2-1996, U.S. Guide to the Expression of Uncertainty in Measurement, National Conference of Standards Laboratories, 1996.

5.    NAMAS NIS 3003, “The Expression of Uncertainty and confidence in Measurement for Calibrations,” Edition 8, NAMAS Executive, National Physical Laboratory, Teddington, TW11 0LW, England, 1995.

6.    Reflectometer/Mismatch Error Limits Calculator, Hewlett-Packard Co., Lit. No. 5952-0948. Write Inquiries Mgr., Hewlett-Packard Co., 5301 Stevens Creek Blvd., Santa Clara, CA 95052.

7.    B. N. Taylor and C. E. Kuyatt, Guidelines for Evaluating and Expressing the Uncertainty of NIST Measurement Results, NIST Technical Note 1297, National Institute of Standards and Technology.

8. NCSL Recommended Practice RP-12, Determining and Reporting Measurement Uncertainties, National Conference of Standards Laboratories.

9. I.A. Harris and F.L. Warner, Re-examination of Mismatch Uncertainty when Measuring Microwave Power and Attenuation, Proceedings of the British IEE, Vol 128, pp 35-41, Feb 1981.

62

VII. Power Measurement Instrumentation Compared

All the previous discussion about power measurement equipment still leaves the important question: which power meter and sensor technology should be used for a particular power measurement? Each method of measuring average power has some advantages over the others, so the answer to that question becomes dependent on the particular measurement situation. Factors such as cost, frequency range, the range of power levels to be measured, the importance of processing and capturing data, accuracy, speed of measurement, and the skill of the personnel involved take on varying degrees of importance in different situations. This chapter compares the measurement systems from several aspects to aid in the decision-making process for any application.

Accuracy vs. Power Level

The first comparison of power measuring systems demonstrates the measurement uncertainty and power range of several equipment selections. The HP EPM series power meters and HP E series sensors were emphasized, although several existing sensors were included. Figure 7-1 shows plots of the RSS uncertainty when measuring power at various levels from - 70 to +20 dBm. The measurement conditions were assumed for a CW signal at 2 GHz and a source SWR of 1.15, and data sheet specifications.

The three parts of this figure are divided to show a comparison of three common combinations of power meter and sensor:

a) HP 432A analog power meter plus HP 8478B thermistor sensor.

b) HP E4418A digital power meter plus exsisting HP 8481A thermocouple and HP 8484D diode sensor.

c) HP E4418A digital power meter plus HP E4412A extended dynamic range power sensor.

The data for Figure 7-1 were computed using a commercially-available mathematics simulation software product called MathCad. To present these operating performances under typical present-day conditions, the ISO uncertainty combining process of Chapter VI was used for the MathCad calculations. Results are approximate, although they are entirely suitable for these comparison purposes.

The reason for presenting these overall measurement uncertainties in this format is that, as far as the user is concerned, there is little need to know whether the sensor works on the diode principle or on the thermocouple principle. And, with the introduction of the new extended-range PDB diode sensors, a single HP E4412A sensor can achieve the- 70 to +20 dBm power range which previously required a combination of diode and thermocouple sensors.

The top graph of Figure 7-1 describes the thermistor sensor/meter combination and is shown mostly for reference. With the decreasing applications of thermistor-type sensors, the primary need for understanding their theory and practice is that they are used as power transfer devices for metrology round robins or for transferring a power reference from a higher-accuracy echelon or national standards labs to operating labs.

A comparison of the top two graphs of Figure 7-1 (A) and (B), shows that the uncertainties of the thermocouple and diode-based systems (B) are somewhat less than the thermistor-based systems (A). At this 2 GHz

63

calculation frequency, the thermocouple and diode sensors have the better SWR (See Figure 7-2), but the thermistor system, being a DC substitution system, does not require a power reference oscillator and its small added uncertainty. These two effects tend to offset each other for this application. The significant advantage of the HP E4418A power meter measurement is the flexibility of being able to use the installed base of all the other HP family of thermocouple and diode sensors.

The third graph of Figure 7-1 (C), for the HP E4418A power meter and HP E4412A extended dynamic range sensor, immediately shows that even with its wide dynamic measurement range from - 70 to +20 dBm, it provides approximately equivalent uncertainties. The dashed portion of the E-series sensor curve (0 to +20 dBm) represents nominal high-power cal factor uncertainty limitations imposed by the sensor, meter, and the calibration system. Refer to the latest sensor technical specifications (literature number 5965-6382E) to determine actual uncertainties for your particular application.

Figure 7-1. RSS uncertainty vs. dynamic power range from data sheet specs for source SWR = 1.15 (r s =0.07) and f = 2 GHz.

(A) Analog thermistor mount system.

(B) HP E4418A digital power meter system using HP 8481D diode and HP 8481A thermocouple sensors.

(C) HP E4418A digital power meter and HP E4412A PDB extended-range sensor. RSS-combining method is the same as used in Chapter VI.

Uncertainty %

|  |  | (A) HP 432A plus |  |  | HP 8478B |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |

-70          -60          -50          -40          -30          -20          -10                    0          +10        +20

Power dBm

Uncertainty %

HP 8481A HP 8481D

|  |  | (B) HP E4418A plus HP 8481D and 8481A |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| M |  |  |  |  |  |  |  |  |  |
| 10 | \ |  |  |  | •. |  |  |  |  |
|  | \ | V |  |  | \ |  |  |  |  |
| b |  | "v |  |  | —!--- |  |  |  |  |

-70          -60          -50          -40          -30          -20          -10                0            +10        +20 Power dBm

(C) HP E4418A plus HP E4412A

Uncertainty %

| \ | i |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | V |  |  |  |  |  | ....... | ....... |

-70          -60          -50          -40          -30          -20          -10                    0          +10        +20

Power dBm

While most modern power meter designs have utilized digital architectures, analog-based meters, such as the HP 432A and HP 435A, are still available. Analog meter measurements are limited by the mechanical meter movement of the instrument which requires uncertainty to be stated in percent of full scale. Thus, at the low end of each range, the uncertainty becomes quite large when expressed as a percent of the reading. Digital systems are free of those problems and, with proper design and an adequate digital display resolution, provide better accuracy.

64

The instrumentation accuracy for a digital meter is specified as a percent of the reading instead of as a percent of full scale. This means that at the point of each range change, there is not a big change in accuracy for the digital meter. This effect can be seen in the max-min excursions of the sawtooth-like curves of the analog meter shown in Figure 7-1 (A). For this reason, the digital power meter does not need as many ranges; each digital range covers 10 dB with little change in accuracy across the range. Thus the HP 437B digital power meter is more accurate than the HP 435A analog power meter.

One application’s advantage attributed to analog meters is for use in “tweaking” functions where an operator must adjust some test component for optimum or maximum power. Digital displays are notoriously difficult to interpret for “maximum” readings, so modern digital meters usually contain a simple “peaker” analog meter. The display of the new HP E4418A power meter features an analog scale in graphic display format which provides that “virtual-peaking” function.

It should be recognized that the accuracy calculations of Figure 7-1 are based on specification values. Such specifications are strongly dependent on the manufacturers’ strategy for setting up their specification budget process. Some published specifications are conservative, some are less so. Manufacturers need to have a good production yield of instruments for the whole family of specifications, so this often leads to a policy of writing specifications which have generous “guard bands” and thus are more conservative.

Further, a particular measurement configuration is likely to be close to one spec limit but meet another spec easily; a second system might reverse the roles. By using the new ISO RSS-uncertainty-combining method, this takes advantage of the random relationship among specifications, and the uncertainties tend to be smaller, yet realistic.

A second reason to observe is that the Figure 7-1 calculations are done for one particular frequency (2 GHz) and one particular source SWR (1.15). A different frequency and different source match would give a different overall uncertainty. Sources frequently have larger reflection coefficient values which would raise the overall uncertainty due to usually-dominant mismatch effects.

Frequency Range and SWR (Reflection Coefficient)

All three types of HP power sensors have models that cover a frequency range from 10 MHz to 18 GHz, some higher, with coaxial inputs. A special version of the HP thermistor mount operates down to 1 MHz (see Chapter III) and the HP 8482A/H thermocouple power sensors operate down to 100 kHz. The effective efficiency at each frequency is correctable with the Calibration Factor dial or keyboard of the power meter, so that parameter is not particularly critical in deciding on a measurement system.

The sensor’s SWR performance is most important because mismatch uncertainty is usually the largest source of error, as described in Chapter VI. Figure 7-2 shows a comparison of the specification limits for the SWR of a thermistor mount, a thermocouple power sensor, an HP 8481D PDB diode power sensor, as well as the HP E series power sensors. It should be recognized that published SWR specifications are usually conservative and that actual performance is often substantially better, yielding lower uncertainty in practice. That fact argues for a measurement process which

65

Figure 7-2. A comparison of specified SWR limits for the HP 8478B thermistor mount, HP 8481A thermocouple power sensor, HP 8481D PDB power sensor, and HP E4412A PDB

measures actual source SWR for situations where highest accuracy is important. These graphs indicate that over the bulk of the frequency range, the thermocouple and diode sensors have a considerably-lower SWR than the thermistor sensor. It also shows that the HP E4412A sensor, even with its superior dynamic range, still provides a satisfactory SWR.

| 1.7 |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1.6 |  |  |  |  |  |  |  | HP 8478B HP 8481D HP 8481A HP E4412A |  |  |  |  |
|  |  |  |  |  |  |  | ------ |  |  |  |  |  |
| 1.5 |  |  |  |  |  |  | ------ |  |  |  |  |  |
|  |  |  |  |  |  |  | HP 8478B Thermistor |  |  |  |  | HP 8481D Diode |
| 1.3 | 3 |  |  |  | HP E4412A Diode |  | i |  |  |  | ,---- |  |
|  |  |  |  |  | HP 8481A Thermocouple |  |  |  |  |  |  |  |
|  |  | . |  |  |  |  |  |  | I 12.4 |  |  |  |
|  |  | 0 | I 50 |  |  |  | 1 I 2 4 |  |  |  |  |  |
|  | 10 MHz 3 | 0 | 50 | 100 | MHz | 1 G | Hz | 24 |  | 10 GHz |  | 30 GHz |

Frequency

Waveguide Sensor Calibration

Power measurements in rectangular waveguide present several special considerations. HP waveguide thermistor sensors (8.2 to 40 GHz) have one advantage. Since thermistor sensors are closed loop there is no need for a 50 MHz power reference oscillator, although this advantage is somewhat offset by their higher SWR performance.

Waveguide thermocouple and diode sensors must have the usual 50 MHz reference oscillator to adjust for calibration factor from one sensor to another. Such a low-frequency signal cannot propagate in a waveguide mode. HP waveguide thermocouple sensors (26.5 to 50.0 GHz) and waveguide diode sensors (26.5 to 50.0 GHz and 75 to 110 GHz) all utilize a special 50 MHz injection port which applies the reference oscillator output to the sensor element in parallel to the usual waveguide input.

Speed of Response

To measure the lowest power ranges with optimum accuracy, power meters are designed with a highly-filtered, narrow bandwidth compared to most other electronic circuits. Narrow band circuits are necessary to pass the desired power-indicating signal but reject the noise that would obscure a weak signal. Narrow bandwidth leads to the long response time. For heat responding power sensors, like the thermistor and thermocouple, response time is also limited by the heating and cooling time constants of the heat

sensor.

66

sensing element. The typical thermistor power measurement has a 35 millisecond time constant and 0 to 99 percent response time of about five time constants or 0.175 seconds. The power meters for thermocouple and PDB sensors have 0 to 99 percent response times of 0.1 to 10 seconds, depending on the range of the power meter. The more sensitive ranges require more averaging and hence longer settling times.

For manual measurements, the speed of response is seldom a problem. By the time the observer turns on the RF power and is ready to take data, the power meter has almost always reached a steady reading.

For analog systems applications, where rapid data acquisition is required, or where the power meter output is being used to control other instruments, the power meter acts like a low pass filter. The equivalent cutoff frequency of the filter has a period roughly the same as the 0 to 99 percent response time. For signals where the power changes too rapidly for the power meter to respond, the power meter averages the changing power. When a power meter is being used to level the output of a signal generator whose frequency is being swept, the speed of the frequency sweep may have to be reduced to allow the power meter time to respond to the power level changes.

There is no clear-cut advantage with regard to speed of one power measurement system over another. In some power ranges one system is faster, and in other ranges another system is faster. If response time is important, manufacturers’ data sheets should be compared for the particular application.

Automated Power Measurement

Recognizing that a large percentage of digital power meters are used in production test and in automated systems, it is reasonable to assume that digitizing measurement speed is critical in at least some of those applications. Digital power meters programmed for automatic operation gather data rapidly and with minimum errors. The data can be processed and analyzed according to programmed instructions, and the system can be operated with little process attention. Even in a manual mode, digital indications are less prone to the human error of misinterpreting the meter scale and using wrong range multipliers. In the case of power measurement, there are additional advantages to automatic systems. Successive data points can be compared mathematically to assure that the power measurement has reached steady state and multiple successive readings can be averaged to statistically reduce the effects of noise.

Measurement speed for data acquisition often becomes a determing factor between competing products, whereas the realizable digitizing speed is usually limited by the response time of the sensor and the need to heavily filter the analog amplified signal on the most sensitive power ranges. For example, the HP 437B specifies a data output of 20 readings per second with a free-running trigger, which simply operates the analog-to-digital circuits at their maximum speed. This recognizes that the analog input to the digitizer is slower in response than the maximum speed of the data output. For example, the HP 437B specifies a 7.0 sec settling time on the most sensitive range, assuming 0.01 dB resolution. This function allows the engineer to take the raw digital data and perform other digital signal processing functions on that data. This method is often accurate enough for particular applications and indeed is faster than waiting for the sensor to respond in an analog sense.

67

The HP E4418A power meter has been optimized for maximum digitizing speed. Since its architecture is totally DSP-based, and it is married to a new HP E series diode sensors, circuit decisions were made to increase the digitizing speed to maximum. For example, output filtering on the sensor is smaller, which provides faster response. On the lower power ranges, this smaller filtering might cause an increase in measurement noise, but the power meter itself provides for digital averaging up to 1,024 readings to minimize noise effects. The meter is specified to provide up to 20 readings per second and 40 per second in the X2 mode. 200 readings per second are specified for the FAST range in the free-run trigger mode, using the binary output format. For that function, circuit settling times are 5 mS for the top 70 dB power ranges.

Susceptibility to Overload

The maximum RF power that may be applied to any power sensor is limited in three ways. The first limit is an average power rating. Too much average power usually causes damage because of excessive accumulated heat. The second limit is the total energy in a pulse. If the pulse power is too high for even a very short time, in spite of the average power being low, the pulses cause a temporary hot spot somewhere in the sensor. Damage occurs before the heat has time to disperse to the rest of the sensor. The third limit is peak envelope power. This limit is usually determined by voltage breakdown phenomena that damages sensor components. The limits are usually stated on the manufacturer’s data sheet. None of the three limits should be exceeded. The power limits of any sensor may be moved upward by adding an attenuator to pre-absorb the bulk of the power. Then the power limits are likely to be dictated by the attenuator characteristics, which, being a passive component, are often fairly rugged and forgiving.

A chart of power limits (Figure 7-3) shows that the HP 8481H power sensor, which consists of a 20-dB attenuator integrated with a thermocouple sensor element, excels in all respects, except for peak envelope power where the thermistor mount is better. One characteristic, that might be important but not obvious from the chart, is the ratio of maximum average power to the largest measurable power. The HP 8481D PDB sensor can absorb 100 mW (+20 dBm) of average power, while the high end of its measurement range is 10 µW (- 20 dBm). This means that the PDB is forgiving in situations where the power level is accidentally set too high. A mistake of 10 dB in setting a source output attenuator, during a measuring routine, will merely cause an

|  | HP 8478B Thermistor Sensor | HP 8481A Thermocouple Sensor | HP 8481H Thermocouple Sensor | HP 8481D Diode Sensor | HP E4412A Extended-range Diode Sensor |
| --- | --- | --- | --- | --- | --- |
| Maximum Average Power | 30 mW | 300 mW | 3.5 W | 100 mW | 200 mW |
| Maximum Energy per Pulse | 10 W • µS | 30 W • µ S | 100 W • µS | (1) | (1) |
| Peak Envelope Power | 200 mW | 15 W | 100 W | 100 mW | 200 mW |

(1) Diode device response is so fast, device cannot average out high-energy pulses.

Figure 7-3. Maximum power limits of various power sensors.

68

off-scale reading for the HP 8481D. The same mistake might damage the other sensors. Excessive power is, by far, the primary cause of power sensor failure.

In most situations the decision about which measurement system to use will probably come to be one of flexibility compared to cost. The flexibility comes in the form of the possibility for automatic measurement and of a large dynamic range of measurement. Accuracy and speed of response are substantially the same in the systems discussed, with the advantage going to a digital power meter. Specifications are likely to change with time, as is cost, so current data sheets should be consulted. The HP EPM power meter series and HP E series sensors are optimized for high speed data acquisition. Yet, with the flexibility offered by the DSP architecture, the meters can easily perform extensive data averaging to minimize noise.

Signal Waveform Effects

While the waveform considerations were fully covered in Chapter V, it is well to consider the waveform factor as a differentiator for the various meters and sensor technology. Briefly, the thermistor is a totally heat-based sensor, and therefore the thermistor sensors handle any input waveform with any arbitrary crest factor, that is, they are true square law sensing elements. Thermocouple sensors are full square law sensing for the same reason, but HP thermocouples operate beyond the thermistor high limit of 10 mW, all the way to 100 mW and 3 watts for the HP 848X H-models which have the integrated fixed pads. The HP 8481B features a 25-watt internal attenuator, and operates from 10 MHz to 18 GHz for medium power applications.

PDB-diode-based sensors of the HP 8481D family feature full square-law performance because their operating power range is limited to a top level of - 20 dBm, thus restricting their meter indications to the square-law range of diodes. The user should assure that peak power excursions do not exceed - 20 dBm.

The HP E series diode sensors require simple attention to their input signal characteristics. CW signals may be applied all the way from - 70 to +20 dBm with confidence and accuracy.

69

VIII Peak Power Instrumentation

A Brief History of Peak Power Measurements

Historically, the development of radar and navigation systems in the late 1930s led to the application of pulsed RF and microwave power. Magnetrons and klystrons were invented to provide the pulsed power, and peak power measurement methods developed concurrently. Since the basic performance of those systems depended primarily on the peak power radiated, it was important to have reliable measurements.1

Early approaches to pulse power measurement have included the following techniques: (1) average power-duty cycle; (2) notch wattmeter; 3) DC-pulse power comparison; (4) barretter integration. Most straightforward is the method of measuring power with a typical averaging sensor, and dividing the result by the duty cycle measured with a video detector and an oscilloscope.

The notch wattmeter method arranged to combine the unknown pulsed signal with another comparison signal usually from a calibrated signal generator. By appropriate video synchronization, the generator signal was “notched out” to zero power at the precise time the unknown RF pulse occurred. A microwave detector responded to the combined power, and allowed the user to set the two power levels to be equal on an oscilloscope trace. The unknown microwave pulse was equal to the known signal generator level, corrected for the signal attenuation in the two paths.

The DC-power comparison method involved calibrating a stable microwave detector with known power levels across its dynamic range, up into its linear detection region. Then, unknown pulsed power could be related to the calibration chart. The early HP 8900A peak power meter was an example of that method.

Finally, barretter integration instrumentation was an innovative solution which depended on measuring the fast temperature rise in a tiny metal wire sensor which absorbed the unknown peak power.2 By determining the slope of the temperature rise in the sensor, the peak power could be measured, the higher the peak, the faster the heat rise and greater the heat slope. The measurement was quite valid and independent of pulse width, but unfortunately, barretters were fragile and lacked great dynamic range. Other peak power meters were offered to industry in the intervening years.

Peak Power Analyzers

Figure 8-1. The HP 8990A peak power analyzer has two 500 MHz to 40 GHz microwave sensor channels and two video channels which provide for comprehensive characterization of RF/microwave pulsed power.

70

Figure 8-2. Typical envelope of pulsed system with overshoot and pulse ringing, shown with 13 pulse parameters which characterize time and amplitude.

As stable PDB diodes came onto the power measurement scene, and combined with powerful data processing technology of the late 1980’s, HP introduced the first complete solution to peak power measurement in 1990. The HP 8990A peak power analyzer represented a comprehensive solution to pulsed power characterization because it was able to measure or compute 13 different parameters of microwave pulsed power.3, 4

Amplitude parameters

1. Pulse-top amplitude

2. Pulse-base amplitude

3. Peak power

4. Overshoot

5. Average power

Time parameters

1. Rise time

2. Fall time

3. Pulse width

4. Off time

5. Duty cycle

6. PRI (pulse repitetion interval)

7. PRF (pulse repetition frequency)

8. Pulse delay

The HP peak power analyzer features dual sensors, with frequency coverage from 500 MHz up to 40 GHz (3 different models), and with a dynamic power range from - 32 to +20 dBm. The dual channels allow the measurement of power ratios, or detection of power at different points in a power transmission or amplification chain. The new GaAs dual-diode sensors were implemented as MMIC components to optimize frequency and temp-erature performance.

Since the balanced diode detectors were operated from the square-law range up through the transition and linear range, each sensor contains an individual EEPROM to store the detection calibration information, and a thermistor sensor to measure and feedback the temperature environment. The PDB sensors are capable of video outputs of 100 MHz bandwidth and thus faithfully preserves the pulse envelope to provide accurate characterization of modern fast rise/fall time pulses.

A three-dimensional calibration scheme is used to completely characterize the sensor performance across its nonlinear detection range vs. frequency response vs temperature. On the production line, sensors are calibrated across a temperature range of 0 to +55° C, and the specified limits of the power and frequency. In operation, the user simply keys in the operating frequency of the unknown input power for lowest uncertainty. As in other HP power meters, a reference power source is available on the front panel with an output of 10 mW at 1.05 GHz, square-wave modulated. This provides the power traceability needed in most production measurement situations.

71

The analyzer itself is essentially a dual channel sampling oscilloscope, fully customized for pulsed power. In addition to the two power channels, there are also two video channels which are primarily used for amplifying triggering signals. But they can also display radar video control signals, and thus provide system time delay measurements. The dual amplification of the power channels is followed by 8-bit ADC and extensive data processing. Random repetitive sampling is used since it can sample at less than the Nyquist rate and still avoid aliasing.

IEEE Video Pulse Standards Adapted for Microwave Pulses

With the envelopes of the two peak power channels amplified and converted to digital form, the processing of such data is limited only by the creativity of the designers. And the digital properties also permit the statistical evaluation of the various pulse parameters. For example, the IEEE standard for determining pulse top characteristics recommends the use of a histogram technique for the pulse top and pulse base, as shown in Figure 8-3.

Figure 8-3. IEEE method for determining pulse top and pulse base with the histogram method. The HP peak power analyzer digital-sampling statistics process makes this measurement using histogram data processing. ANSI/IEEE

Std 181-1977,

©

Copyright          1977,

IEEE, all rights reserved.

One reason that pulsed power is more difficult to measure is that waveform envelopes under test may need many different parameters to characterize the power flow as shown in Figure 8-2. Interestingly, some parameters are measured directly, while others are computed from those direct measurements. Further, while industry-accepted terminology for pulsed RF carriers have been in common use for decades for a few of the key parameters, no nationally-written standard was available. HP chose to adapt the IEEE standards for video pulse characterization and measurement. These are shown in Figure 8-4, which adapts some IEEE definitions to the HP waveform envelope terminology. Two standards are applicable:

1) ANSI/IEEE STD 181-1977, “IEEE Standard on Pulse Measurement and Analysis by Objective Techniques,” July 22, 1977. (Revised from 181-1955, “Methods of Measurement of Pulse Qualtities.” [5]

2) IEEE STD 194-1977, “IEEE Standard Pulse Terms and Definitions,” July 26, 1977. [6]

72

Figure 8-4. IEEE pulse definitions and standards for video parameters applied to micro-wave pulse envelopes. ANSI/IEEE Std 194-

©

1977, Copyright

1977, IEEE all rights reserved.

It was recognized that while terms and graphics from both those standards were written for video pulse characteristics, most of the measurement theory and intent of the definitions can be applied to the waveform envelopes of pulse-modulated RF and microwave carriers. Several obvious exceptions would be parameters such as pre-shoot, which is the negative-going undershoot that precedes a pulse risetime. Negative power would be meaningless. The same reasoning would apply to the undershoot following the fall time of a pulse.

For measurements of pulse parameters such as risetime or overshoot to be meaningful, the points on the waveform that are used in the measurement must be defined unambiguously. Since all the time parameters are measured between specific amplitude points on the pulse, and since all the amplitude points are referenced to the two levels named “top” and “base,” Figure 8-4 shows how they are defined.

Peak Power Waveform Definitions

The following are the definitions for the 13 RF pulse parameters as adapted by HP from IEEE video definitions:

Rise time The time difference between the proximal and distal first transition points, usually 10 and 90 percent of pulse-top amplitude (vertical display is linear power).

Fall time      Same as risetime measured on the last transition.

Pulse width                  The pulse duration measured at the mesial level;

normally taken as the 50% power level.

Off time        Measured on the mesial (50%) power line; pulse separation, the interval between the pulse stop time of a first pulse waveform and the pulse start time of the immediately following pulse waveform in a pulse train.

Duty cycle                      The previously measured pulse duration divided by the

pulse repetition period.

73

PRI                        (Pulse Repetition Interval) The interval between the pulse

start time of a first pulse waveform and the pulse start time of the immediately following pulse waveform in a periodic pulse train.

PRF                      (Pulse Repetition Frequency) The reciprocal of PRI.

Pulse delay                    The occurance in time of one pulse waveform before (after)

another pulse waveform; usually the reference time would be a video system timing or clock pulse.

Pulse-top                          Pulse amplitude, defined as the algebraic difference

amplitude                        between the top magnitude and the base magnitude;

calls for a specific procedure or algorithm, such as

the histogram method.

Pulse-base                      The pulse waveform baseline specified to be obtained by

amplitude                        the histogram algorithm.

Peak power                  The highest point of power in the waveform, usually at

the first overshoot; it might also occur elsewhere across the pulse top if parasitic oscillations or large amplitude ringing occurs; peak power is not the pulse-top amplitude which is the primary measurement of pulse amplitude.

Overshoot A distortion that follows a major transition; the difference

between the peak power point and the pulse-top amplitude computed as a percentage of the pulse-top amplitude.

Average          Computed by using the statistical data from amplitude            power and time measurements; should have been called pulse-average power.

Measuring Complex Waveforms other than Pulsed Power

The peak power analyzer can also make meaningful measurements on non-pulsed signals which have high-data-rate or other complex modulation formats. For example, setting infinite persistence on the display, the amplitude transitions can be recorded for a 16QAM (quadrature-amplitude-modulated) communications signal. The constellation diagram of 16QAM has three circles of constant amplitude, and system designers often wish to examine the compression effects of their components by looking at the amplitude of the corner states of the constellation diagram. Using power markers of the analyzer, these relative compression measurements can easily be made, while an average power meter would be useless for such discrimination.

The peak power analyzer can also be used for pulsed component tests, such as amplifiers. By using the ratio process and two sensors, one at the input and one at the output of the pulsed component under test, the 1 dB compression point may be determined. As the input power is increased, the analyzer monitors the component gain, and when the reference gain value drops by 1 dB, that indicates the point of compression desired.

Complex signals with modulation bandwidths within the 100 MHz bandwidth of the HP 8990A and 8992A sensors can be measured and displayed. For example, two-tone tests which are used to determine intermodulation can be characterized as long as the separation of the two tones is less than 100 MHz.

74

The HP 8992A digital video power analyzer is customized for digital transmission applications. Its ability to detect random peak power events make it possible to monitor receiver headroom and digital modulation quality.

1. R.E. Henning, “Peak Power Measurement Technique,” Sperry Engineering Review, (May-June 1955).

2. M. Skolnik, “Introduction to Radar Systems,” McGraw-Hill, Inc., (1962).

3. D. Scherer, “Designing Sensors to Read Peak Power of Pulsed Waveforms,” Microwaves & RF, February, 1990.

4. D. Scherer, et al, “The Peak Power Analyzer, a New Microwave Tool,” HP Journal, (April, 1992).

5.      ANSI/IEEE STD181-1977, “IEEE Standard on Pulse Measurement and Analysis by Objective Techniques,” July 22, 1977. Revised from 181-1955, “Methods of Measurement of Pulse Qualtities.”

6. IEEE STD 194-1977, “IEEE Standard Pulse Terms and Definitions,” (July 26, 1977).

75

Glossary and List of Symbols

ADC

a g

ANSI

AM

a ,

b g

b ,

b s

C b

C c

C o

CW

C, C , C 12

dB

dBm

D

DSP

e

emf

e p

E rms

e s

f m

f r

FET

FM

GaAs

i

i

i p

I rms

I s

ISO

K

K b

Kc

L

L w

m

m i

MMIC

Mu

Muc

n

N

NAMAS

NBS

NCSL

NIST

NPL

P

P

P av

P avg

P cal

P d

P fs

P

analog-digital converter

incident wave upon a generator

American National Standards Institute

amplitude modulation

incident wave upon a load

emerging wave from a generator

reflected wave from a load

generated wave of a source

bypass capacitance

coupling capacitance

diode junction capacitance

continuous wave

capacitors

decibel

decibels referenced to 1 mW

power meter drift

digital signal processor

instantaneous voltage

electromotive force

peak voltage

root mean-square of a voltage waveform

source voltage

maximum modulation frequency component

repetition frequency

field effect transistor

frequency modulation

Gallium arsenide

instantaneous current

instrumentation uncertainty

load current

peak current

root mean-square of a current waveform

diode saturation current

International Standards Organization

Boltzmann’s constant

calibration factor

sensor cal factor at cal frequency

inductance

wire lead inductance

power meter magnification (gain)

instrument magnification uncertainty

microwave monolithic integrated circuit

gain due to mismatch between unknown generator and sensor

gain due to mismatch between sensor and cal source

a diode correction constant

power meter noise

National Measurement Accreditation Scheme (UK)

National Bureau of Standards (now NIST)

National Conference of Standards Laboratories

U.S. National Institute of Standards & Technology (formerly NBS)

National Physics Laboratory

product of voltage and current

power

available generator power

average power

power delivered to Z      load by meter cal source o

dissipated power

power at full scale

net power transferred to load from generator

76

| P gZo | power delivered to Z load from generator o |
| --- | --- |
| P i | incident power |
| P m | meter indication of power |
| Pmc | power level indicated during calibration |
| P p | pulse power |
| P r | reflected power |
| P ref | reference power |
| P rf | radio frequency power |
| P sub | substituted power, DC or low frequency equivalent of an RF power |
| PDB | planar-doped-barrier (diode) |
| PTB | Physikalisch-Technische Bundesanstalt |
| P l | power sensor linearity |
| q | charge of electron |
| R | resistance |
| RF | radio frequency |
| RSS | root-sum-of-the-squares |
| Re | resistance of compensating thermistor |
| R d | resistance of detecting thermistor |
| R0 | diode origin resistance |
| R, Rj, R 2 , R l | resistor |
| R b | bulk resistance of silicon |
| RR | round robin |
| R s | source resistance |
| R T | thermistor resistance |
| t | time as a variable |
| t | power meter translation (offset) error |
| T | temperature in Kelvins |
| T | time lapse |
| T o | period of a waveform |
| T , | period of the lowest frequency |
| t r | period of the repetition frequency |
| SWR 1 | voltage standing wave ratio |
| SI | International System of Units |
| u | standard uncertainty |
| U | expanded uncertainty (for example catalog spec) |
| v | instantaneous voltage |
| v | voltage across a load |
| v o | output voltage |
| V 0 , V 1 , V 2 , V T | Voltages |
| V c | Voltage driving the compensating bridge |
| V h | Peltier emf at a hot junction |
| V rf | voltage driving the rf thermistor bridge |
| V rfo | voltage driving the rf thermistor bridge when no rf power is applied |
| W | watt |
| Zc | power meter zero carryover value |
| Zr | reference impedance |
| Zs | power meter zero set value |
| Z o | reference impedance |
| Z | generator impedance |
| Z | load impedance |
| a | q/nKT |
| g | complex reflection coefficient looking back into a generator |
| , | complex reflection coefficient of a load |
| ^e | effective efficiency |
| p | reflection coefficient magnitude of a load |
| Pg | reflection coefficient magnitude of a generator |
| T | pulse width |
| <l | phase angle between a sinusoidal waveform and a reference waveform |
| 9g | reflection coefficient angle of a generator |
| (¦, | reflection coefficient angle of a load |
| O | ohms |

1. Due to infrequent use of the term power standing wave ratio, common usage in the U.S.A. has shortened VSWR to SWR. Some parts of the world continue to use VSWR to refer to voltage standing wave ratio.

For more information about Hewlett-Packard test and measurement products, applications, services, and for a current sales office listing, visit our web site, http://www.hp.com/go/tmdir. You can also contact one of the following centers and ask for a test and measurement sales representative.

United States:

Hewlett-Packard Company

Test and Measurement Call Center

P.O. Box 4026

Englewood, CO 80155-4026

1 800 452 4844

Canada:

Hewlett-Packard Canada Ltd. 5150 Spectrum Way Mississauga, Ontario L4W 5G1 (905) 206 4725

Europe:

Hewlett-Packard

European Marketing Centre

P.O. Box 999

1180 AZ Amstelveen

The Netherlands

(31 20) 547 9900

Japan:

Hewlett-Packard Japan Ltd. Measurement Assistance Center 9-1, Takakura-Cho, Hachioji-Shi, Tokyo 192, Japan Tel: (81-426) 56-7832 Fax: (81-426) 56-7840

Latin America:

Hewlett-Packard

Latin American Region Headquarters 5200 Blue Lagoon Drive, 9th Floor Miami, Florida 33126, U.S.A. (305) 267 4245/4220

Australia/New Zealand:

Hewlett-Packard Australia Ltd. 31-41 Joseph Street Blackburn, Victoria 3130, Australia 1 800 629 485

Asia Pacific:

Hewlett-Packard Asia Pacific Ltd.

17-21/F Shell Tower, Times Square,

1 Matheson Street, Causeway Bay,

Hong Kong

Tel: (852) 2599 7777

Fax: (852) 2506 9285

Data Subject to Change Copyright © 1998 Hewlett-Packard Company Printed in U.S.A. 6/98 5965-6630E